## ClaimIndex
ClaimIndex — every public claim, its evidence and status | [6 empty] | About / navigation →
A source match is not clinical validation. Status counts below are live COUNTIF formulas. Workbook value = formula over Sources; Check = stated value within half a unit of its last decimal of the workbook value.
| Status | Claims
| supported |  | The stated value or fact matches a saved artifact, report or dataset record named in Sources.
| supported-arithmetic |  | Follows by arithmetic from stated definitions (formula shown); no experiment involved.
| tested |  | Behaviour covered by a test that passed in the saved test record (JUnit) named in the row.
| implemented |  | Present in code or configuration only; no saved passing test is cited for the behaviour.
| negative-result |  | A saved study result that did not meet its declared bar; reported as negative, not as a gain.
| author-reported |  | Stated by the team without a saved log, artifact or primary source in the repository.
| proposed/illustrative |  | A proposal, scenario or illustration; not an evidential claim.
| pending |  | Evidence is not on this installation or a primary-source check is still to be done.
| unsupported/contradicted |  | The saved source disagrees with the stated value, or no support exists.
total
Claim ID | Exact claim | Location(s) | Evaluation / model scope | Status | Source IDs | Workbook value (formula) | Stated value | Decimals | Check | Material limitation | Unresolved action | Manifest group ¦ check
C-SRV-01 | The app serves lightgbm.dyn-prof-v2.gt-primary-2026-10-13-fusion for recorded data | README.md (Results at a glance); reports/evidence_manifest.md; README §7 | Software configuration | supported | reg.serving; registry | — | — |  | — |  |  | serving-identity ¦ serving model (recorded data)
C-FC-01 | Serving MAE 12.96 mg/dL at 30 min | README.md (Results at a glance); reports/evidence_manifest.md; Deck s11; Deck s13; narration scene 8 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.serv.mae30 |  | 12.96 | 2 |  | Same 14 people as training; macro over participants |  | demo-forecast ¦ MAE 30
C-FC-02 | Serving MAE 20.26 mg/dL at 60 min | README.md (Results at a glance); reports/evidence_manifest.md; Deck s11; Deck s13 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.serv.mae60 |  | 20.26 | 2 |  | Same 14 people as training |  | demo-forecast ¦ MAE 60
C-FC-03 | Persistence MAE 14.84 mg/dL at 30 min (same origins) | README.md (Results at a glance); reports/evidence_manifest.md; Deck s11; Deck s13 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.pers.mae30 |  | 14.84 | 2 |  |  |  | demo-forecast ¦ persistence 30
C-FC-04 | Persistence MAE 23.23 mg/dL at 60 min (same origins) | README.md (Results at a glance); reports/evidence_manifest.md; Deck s11; Deck s13 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.pers.mae60 |  | 23.23 | 2 |  |  |  | demo-forecast ¦ persistence 60
C-WN-01 | 61 new high-glucose onsets warned (60-min warning) | README.md (Results at a glance); reports/evidence_manifest.md; Deck s13 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.ev.detected_onsets |  | 61 | 0 |  | Research event definition; not triage |  | demo-warning ¦ onsets warned
C-WN-02 | 64 new high-glucose onsets in the final test segment | README.md (Results at a glance); reports/evidence_manifest.md; Deck s13 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.ev.true_onsets |  | 64 | 0 |  | Onsets ≠ origins ≠ participants |  | demo-warning ¦ onsets
C-WN-03 | 40 false alarms | README.md (Results at a glance); reports/evidence_manifest.md | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.ev.false_alarm_episodes |  | 40 | 0 |  |  |  | demo-warning ¦ false alarms
C-WN-04 | 1.49 false alarms per observed patient-day | README.md (Results at a glance); reports/evidence_manifest.md; Deck s13 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.ev.false_alarm_episodes; fus.ev.observed_patient_days |  | 1.49 | 2 |  | Denominator = observed patient-days from saved event metadata (approximate time base: issued origins × 5 min) |  | demo-warning ¦ false alarms per observed day
C-WN-05 | Median forecast lead 30 min among detected onsets | README.md (Results at a glance); reports/evidence_manifest.md; narration metric card | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.ev.qualified_lead_median |  | 30 | 0 |  | Denominator = detected onsets with a qualified lead (61), not all onsets |  | demo-warning ¦ median forecast lead
C-WN-06 | Warning active at 11.8 % of eligible times | README.md (Results at a glance); reports/evidence_manifest.md; reports/warning_time_share.md; narration metric card | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | wu.firing; wu.eligible |  | 11.8 | 1 |  | Post-hoc diagnostic |  | demo-time-in-warning ¦ share of eligible times warning
C-RG-01 | Served 80 % ranges covered 74.2 % at 30 min (under-coverage) | README.md (Results at a glance); reports/evidence_manifest.md; Deck s10; Deck s11; Deck s14 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.iv.coverage30 |  | 74.2 | 1 |  | Pooled over origins; per-person coverage varies |  | demo-ranges ¦ demo coverage 30
C-RG-02 | Served 80 % ranges covered 74.0 % at 60 min (under-coverage) | README.md (Results at a glance); reports/evidence_manifest.md; Deck s11; Deck s14 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.iv.coverage60 |  | 74 | 1 |  |  |  | demo-ranges ¦ demo coverage 60
C-RG-03 | Research procedure ranges covered 80.1 % at 30 min (people held out) | README.md (Results at a glance); reports/evidence_manifest.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | nest.cov30 |  | 80.1 | 1 |  | Different procedure from the served model |  | demo-ranges ¦ nested research coverage 30
C-RG-04 | Research procedure ranges covered 78.8 % at 60 min (people held out) | README.md (Results at a glance); reports/evidence_manifest.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | nest.cov60 |  | 78.8 | 1 |  |  |  | demo-ranges ¦ nested research coverage 60
C-NS-01 | Research candidate MAE 13.88 at 30 min (D-032) | README.md (Results at a glance); reports/evidence_manifest.md; Deck s11; narration scene 8 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | nest.sel.mae30 |  | 13.88 | 2 |  | Not served; cohort previously inspected |  | nested-forecast ¦ MAE 30
C-NS-02 | Research candidate MAE 22.30 at 60 min (D-032) | README.md (Results at a glance); reports/evidence_manifest.md; Deck s11 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | nest.sel.mae60 |  | 22.3 | 2 |  | Source value 22.305 is an exact rounding tie at the source's own precision; stated 22.3 is one of its two valid roundings |  | nested-forecast ¦ MAE 60
C-NS-03 | Nested persistence 15.34 at 30 min | README.md (Results at a glance); reports/evidence_manifest.md; Deck s11 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | nest.pers.mae30 |  | 15.34 | 2 |  |  |  | nested-forecast ¦ persistence 30
C-NS-04 | Nested persistence 24.47 at 60 min | README.md (Results at a glance); reports/evidence_manifest.md; Deck s11 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | nest.pers.mae60 |  | 24.47 | 2 |  |  |  | nested-forecast ¦ persistence 60
C-NS-05 | Ridge on CGM only chosen in 4 of 5 folds | README.md (Results at a glance); reports/evidence_manifest.md; Deck s11 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | nest.ridge_folds |  | 4 | 0 |  |  |  | nested-forecast ¦ folds choosing Ridge CGM-only
C-EX-01 | External 15-min LightGBM MAE 11.15 at 30 min | README.md (Results at a glance); reports/evidence_manifest.md | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | ext.cand.mae30 |  | 11.15 | 2 |  | CGM only; tests the forecasting core, not the fusion |  | external-15min ¦ MAE 30
C-EX-02 | External 15-min LightGBM MAE 20.11 at 60 min | README.md (Results at a glance); reports/evidence_manifest.md | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | ext.cand.mae60 |  | 20.11 | 2 |  | Source value 20.105 is an exact rounding tie at the source's own precision; stated 20.11 is one of its two valid roundings |  | external-15min ¦ MAE 60
C-EX-03 | External persistence 12.14 at 30 min | README.md (Results at a glance); reports/evidence_manifest.md | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | ext.pers.mae30 |  | 12.14 | 2 |  | Source value 12.145 is an exact rounding tie at the source's own precision; stated 12.14 is one of its two valid roundings |  | external-15min ¦ persistence 30
C-EX-04 | External persistence 20.44 at 60 min | README.md (Results at a glance); reports/evidence_manifest.md | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | ext.pers.mae60 |  | 20.44 | 2 |  |  |  | external-15min ¦ persistence 60
C-EX-05 | External 60-minute difference is inconclusive (95 % CI upper bound above 0) | README.md (Results at a glance); reports/evidence_manifest.md | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | ext.cmp.hi60; ext.cmp.lo60; ext.cmp.hi60 |  | — |  | see limitation | Interval crosses zero |  | external-15min ¦ 60-minute difference
C-PF-01 | Profile minus no profile (held out) +0.34 mg/dL at 30 min | README.md (Results at a glance); reports/evidence_manifest.md; README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | negative-result | nest.prof.d30 |  | 0.34 | 2 |  | CI [−0.28, +0.89] crosses zero: no reliable benefit; Source value 0.335 is an exact rounding tie at the source's own precision; stated 0.34 is one of its two valid roundings |  | profile-ablation-nested ¦ profile minus no profile, 30 min
C-PF-02 | Profile minus no profile (held out) −0.75 mg/dL at 60 min | README.md (Results at a glance); reports/evidence_manifest.md; README §7; Deck s14 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | negative-result | nest.prof.d60 |  | -0.75 | 2 |  | CI [−2.19, +0.61] crosses zero |  | profile-ablation-nested ¦ profile minus no profile, 60 min
C-PF-03 | Profile ablation 30-min CI lower −0.28 | README.md (Results at a glance); reports/evidence_manifest.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | nest.prof.lo30 |  | -0.28 | 2 |  | Source value -0.275 is an exact rounding tie at the source's own precision; stated -0.28 is one of its two valid roundings
C-PF-04 | Profile ablation 30-min CI upper +0.89 | README.md (Results at a glance); reports/evidence_manifest.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | nest.prof.hi30 |  | 0.89 | 2
C-PF-05 | Profile ablation 60-min CI lower −2.19 | README.md (Results at a glance); reports/evidence_manifest.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | nest.prof.lo60 |  | -2.19 | 2
C-PF-06 | Profile ablation 60-min CI upper +0.61 | README.md (Results at a glance); reports/evidence_manifest.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | nest.prof.hi60 |  | 0.61 | 2
C-2H-01 | Two hours ahead: MAE 30.6 at 120 min (D-043) | README.md (Results at a glance); reports/evidence_manifest.md; narration scene 16 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | 2h.sel.mae120 |  | 30.6 | 1 |  | Research procedure, not served |  | two-hour ¦ MAE 120
C-2H-02 | Two-hour persistence 35.9 at 120 min | README.md (Results at a glance); reports/evidence_manifest.md; narration scene 16 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | 2h.pers.mae120 |  | 35.9 | 1 |  |  |  | two-hour ¦ persistence 120
C-2H-03 | Two-hour procedure better than persistence for 14 of 14 people | README.md (Results at a glance); reports/evidence_manifest.md; narration scene 16 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | 2h.cmp.better |  | 14 | 0 |  |  |  | two-hour ¦ people better
C-2H-04 | Two-hour warning: 315 of 329 onsets | README.md (Results at a glance); reports/evidence_manifest.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | 2h.w.detected_onsets |  | 315 | 0 |  |  |  | two-hour ¦ onsets warned
C-2H-05 | Two-hour warning onsets (denominator) 329 | README.md (Results at a glance); reports/evidence_manifest.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | 2h.w.true_onsets |  | 329 | 0
C-2H-06 | Two-hour warning 1.62 false alarms per observed day | README.md (Results at a glance); reports/evidence_manifest.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | 2h.w.false_alarms_per_observed_day |  | 1.62 | 2 |  |  |  | two-hour ¦ false alarms per observed day
C-2H-07 | Two-hour warning median lead 15 min | README.md (Results at a glance); reports/evidence_manifest.md; narration scene 16 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | 2h.w.qualified_lead_median |  | 15 | 0 |  | Median among detected onsets |  | two-hour ¦ median lead (IQR)
C-2H-08 | Two-hour lead IQR 5–65 min (lower) | README.md (Results at a glance); reports/evidence_manifest.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | 2h.w.qualified_lead_iqr_lo |  | 5 | 0
C-2H-09 | Two-hour lead IQR 5–65 min (upper) | README.md (Results at a glance); reports/evidence_manifest.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | 2h.w.qualified_lead_iqr_hi |  | 65 | 0
C-2H-10 | Two-hour ranges covered 76.5 % (target 80 %) | README.md (Results at a glance); reports/evidence_manifest.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | 2h.cov120 |  | 76.5 | 1 |  |  |  | two-hour ¦ coverage 120
C-2H-11 | Two-hour profile ablation −0.86 [−2.60, +1.02] (LightGBM) | README §7; reports/two_hour_benchmark.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | negative-result | 2h.prof.d |  | -0.86 | 2 |  | CI crosses zero |  | two-hour ¦ profile ablation 120 (LightGBM)
C-2X-01 | External two-hour MAE 29.55 at 120 min (D-045) | README.md (Results at a glance); reports/evidence_manifest.md | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | 2hx.mae120 |  | 29.55 | 2 |  | Second opening of this cohort |  | two-hour-external ¦ MAE 120
C-2X-02 | External two-hour persistence 31.68 | README.md (Results at a glance); reports/evidence_manifest.md | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | 2hx.pers120 |  | 31.68 | 2 |  | Source value 31.675 is an exact rounding tie at the source's own precision; stated 31.68 is one of its two valid roundings |  | two-hour-external ¦ persistence 120
C-2X-03 | External two-hour difference −2.12 mg/dL | README.md (Results at a glance); reports/evidence_manifest.md | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | 2hx.cmp.mean_diff |  | -2.12 | 2 |  |  |  | two-hour-external ¦ difference (95 % CI)
C-2X-04 | External two-hour CI lower −2.67 | README.md (Results at a glance); reports/evidence_manifest.md | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | 2hx.cmp.lo |  | -2.67 | 2
C-2X-05 | External two-hour CI upper −1.53 | README.md (Results at a glance); reports/evidence_manifest.md | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | 2hx.cmp.hi |  | -1.53 | 2
C-2X-06 | Better for 78 of 98 people (two-hour external) | README.md (Results at a glance); reports/evidence_manifest.md | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | 2hx.cmp.improved |  | 78 | 0 |  |  |  | two-hour-external ¦ people better
C-2X-07 | External two-hour 98 participants | README.md (Results at a glance); reports/evidence_manifest.md; Deck s11 | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | 2hx.cmp.n |  | 98 | 0
C-2X-08 | External two-hour MAE 10.90 at 30 min | README.md (Results at a glance); reports/evidence_manifest.md; Deck s11; narration scene 8 | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | 2hx.mae30 |  | 10.9 | 2 |  |  |  | two-hour-external ¦ MAE 30 / 60
C-2X-09 | External two-hour MAE 19.31 at 60 min | README.md (Results at a glance); reports/evidence_manifest.md; Deck s11 | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | 2hx.mae60 |  | 19.31 | 2
C-2X-10 | External two-hour persistence 12.14 at 30 min | README.md (Results at a glance); reports/evidence_manifest.md; Deck s11 | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | 2hx.pers30 |  | 12.14 | 2 |  | Source value 12.145 is an exact rounding tie at the source's own precision; stated 12.14 is one of its two valid roundings |  | two-hour-external ¦ persistence 30 / 60
C-2X-11 | External two-hour persistence 20.44 at 60 min | README.md (Results at a glance); reports/evidence_manifest.md; Deck s11 | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | 2hx.pers60 |  | 20.44 | 2
C-2X-12 | External two-hour ranges covered 87 % | README.md (Results at a glance); reports/evidence_manifest.md | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | 2hx.cov120 |  | 87 | 0 |  | Wider than needed (over-coverage) |  | two-hour-external ¦ coverage 120
C-CL-01 | Classifier warned 329 of 329 onsets (D-044) | README §7; reports/two_hour_classifier.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | negative-result | clf.detected_onsets |  | 329 | 0 |  | Warned at ~99 % of eligible time: degenerate |  | classifier-negative ¦ onsets warned
C-CL-02 | Classifier 0.18 false alarms per day | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | negative-result | clf.false_alarms_per_observed_day |  | 0.18 | 2 |  |  |  | classifier-negative ¦ false alarms per day
C-CL-03 | Classifier warning active at 99 % of eligible time | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | negative-result | clf.share_of_eligible_time_warning |  | 99 | 0 |  |  |  | classifier-negative ¦ share of eligible time warning
C-CL-04 | Classifier ROC AUC 0.64 | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | negative-result | clf.origin_roc_auc |  | 0.64 | 2 |  |  |  | classifier-negative ¦ ROC AUC
C-46-01 | Two-hour coverage before pooled calibration 76.5 % | README.md (Results at a glance); reports/evidence_manifest.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) — re-analysis | supported | wr.b0.cov120 |  | 76.5 | 1 |  |  |  | d046-ranges ¦ coverage 120 before
C-46-02 | Pooled calibration brought two-hour coverage to 79.8 % (D-046 B) | README.md (Results at a glance); reports/evidence_manifest.md; narration scene 16 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) — re-analysis | supported | wr.b1.cov120 |  | 79.8 | 1 |  | Not applied to the demo model |  | d046-ranges ¦ coverage 120 after
C-46-03 | 60-min coverage 78.2 → 79.5 % (before) | README.md (Results at a glance); reports/evidence_manifest.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | wr.b0.cov60 |  | 78.2 | 1 |  |  |  | d046-ranges ¦ coverage 60 before → after
C-46-04 | 60-min coverage 78.2 → 79.5 % (after) | README.md (Results at a glance); reports/evidence_manifest.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | wr.b1.cov60 |  | 79.5 | 1
C-46-05 | Pooled ranges 7 mg/dL wider at 120 min | README.md (Results at a glance); reports/evidence_manifest.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | wr.b1.width120; wr.b0.width120 |  | 7 | 0 |  |  |  | d046-ranges ¦ width increase
C-46-06 | Per-person median coverage 81.2 % at 120 min | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | wr.b1.pmed120 |  | 81.2 | 1 |  |  |  | d046-ranges ¦ per-person median coverage 120
C-46-07 | Earlier warning: 126 vs 123 of 329 warned ≥ 30 min (selected vs incumbent) | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | negative-result | wr.early.sel |  | 126 | 0 |  |  |  | d046-earlier-warning ¦ warned ≥ 30 min ahead, selected vs incumbent
C-46-08 | Incumbent warned 123 of 329 ≥ 30 min ahead | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | wr.early.inc |  | 123 | 0
C-46-09 | Earlier-warning difference +0.9 points | README §7; docs/decisions.md D-046 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | negative-result | wr.diff.d |  | 0.9 | 1 |  | CI [−5.5, +6.6] crosses zero |  | d046-earlier-warning ¦ difference (95 % CI)
C-46-10 | Earlier-warning CI lower −5.5 points | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | wr.diff.lo |  | -5.5 | 1
C-46-11 | Earlier-warning CI upper +6.6 points | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | wr.diff.hi |  | 6.6 | 1
C-46-12 | Median lead still 15 min (D-046 A) | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | wr.lead.sel |  | 15 | 0 |  |  |  | d046-earlier-warning ¦ median lead
C-46-13 | Profile prior −0.53 mg/dL at 120 min (D-046 C) | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | negative-result | wr.prior.all.d |  | -0.53 | 2 |  | CI [−1.57, +0.58] crosses zero |  | d046-profile-prior ¦ 120 min
C-46-14 | Profile prior first 24 h −0.93 mg/dL | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | negative-result | wr.prior.d1.d |  | -0.93 | 2 |  | CI [−2.74, +0.48] |  | d046-profile-prior ¦ 120 min, first 24 h
C-SC-01 | 145 participants imported | README §7 | Data inventory | supported | inv.cgm; inv.sh |  | 145 | 0 |  |  |  | scale ¦ imported
C-SC-02 | 14 evaluated internally | README §7 | Data inventory | supported | inv.cgm_eval |  | 14 | 0 |  |  |  | scale ¦ evaluated internally
C-SC-03 | 98 evaluated externally | README §7 | Data inventory | supported | inv.sh_eval |  | 98 | 0 |  |  |  | scale ¦ evaluated externally
C-50-01 | D-050 candidate pooled ranges 77.9 % at 30 min (test, retrospective) | README §7; reports/serving_ranges_study.md | Serving pipeline, final test (inspected) — candidate, not served | supported | rg.pool.cov30 |  | 77.9 | 1 |  | Research on inspected data; not the served interval |  | ranges-d050 ¦ candidate coverage 30 / 60 (G on validation + calibration, test, retrospective)
C-50-02 | D-050 candidate pooled ranges 78.7 % at 60 min | README §7 | Candidate, not served | supported | rg.pool.cov60 |  | 78.7 | 1
C-50-03 | D-050 served coverage 74.2 % at 30 min | README §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | rg.inc.cov30 |  | 74.2 | 1 |  |  |  | ranges-d050 ¦ served coverage 30 / 60
C-50-04 | D-050 served coverage 74.0 % at 60 min | README §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | rg.inc.cov60 |  | 74 | 1
C-50-05 | D-050: served ranges stay (trend-stratified shape not adopted) | README §7; docs/decisions.md D-050 | Gate | negative-result | Gates: D-050; rg.chosen |  | FAILED |  |  | Pre-test coverage-error gain below the 1.0-point bar |  | ranges-d050 ¦ decision
C-51-01 | D-051 C0 60-min MAE 22.27 | README §7; reports/tune_study.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | tu.C0.mae60 |  | 22.27 | 2 |  |  |  | tune-d051 ¦ C0 / C1 60-min MAE
C-51-02 | D-051 C1 60-min MAE 21.90 | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | tu.C1.mae60 |  | 21.9 | 2
C-51-03 | D-051 C1 − C0 −0.37 mg/dL (small internal gain vs equally tuned C0) | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | negative-result | tu.c1c0.mean_diff |  | -0.37 | 2 |  | Does not reach the declared −0.5 bar; not a release |  | tune-d051 ¦ C1 − C0 (internal)
C-51-04 | D-051 C1 − C0 CI [−0.63, −0.06] (lower) | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | tu.c1c0.lo |  | -0.63 | 2
C-51-05 | D-051 C1 − C0 CI [−0.63, −0.06] (upper) | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | tu.c1c0.hi |  | -0.06 | 2
C-51-06 | D-051 10 of 14 people better | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | tu.c1c0.improved |  | 10 | 0 |  |  |  | tune-d051 ¦ people better
C-51-07 | D-051 C1 − refitted serving recipe −0.65 [−1.41, +0.07] (not established) | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | negative-result | tu.c1r0.mean_diff |  | -0.65 | 2 |  | Different comparator from C0; CI crosses zero |  | tune-d051 ¦ C1 − refitted serving recipe (not established)
C-51-08 | D-051 not adopted | README §7; docs/decisions.md D-051 | Gate | negative-result | Gates: D-051; tu.adopt.c1_better |  | FAILED |  |  |  |  | tune-d051 ¦ adoption
C-52-01 | D-052 S 60-min personalised MAE 21.62 | reports/joint_study.md; docs/decisions.md D-052 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | jt.S.mae60 |  | 21.62 | 2 |  |  |  | joint-d052-forecast ¦ S 60-min personalised MAE
C-52-02 | D-052 R0 60-min personalised MAE 22.50 (refitted serving recipe) | reports/joint_study.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | jt.R0.mae60 |  | 22.5 | 2 |  |  |  | joint-d052-forecast ¦ R0 60-min personalised MAE (refitted serving recipe)
C-52-03 | D-052 S − R0 −0.88 [−1.75, −0.13] at 60 min | reports/joint_study.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | jt.gate.mean_diff_60 |  | -0.88 | 2 |  |  |  | joint-d052-forecast ¦ S − R0 at 60 min (participant bootstrap)
C-52-04 | D-052 primary gate not met (high-target +1.70 vs +1.0) | docs/decisions.md D-052 | Gate | negative-result | jt.gate.high_worsening_60 |  | 1.7 | 2 |  |  |  | joint-d052-forecast ¦ primary gate
C-52-05 | D-052 release guardrails not met | docs/decisions.md D-052 | Gate | negative-result | Gates: D-052 |  | FAILED |  |  |  |  | joint-d052-operational ¦ release guardrails vs R0
C-52-06 | D-052 candidate 80 % range coverage 78.4 % at 60 min | docs/decisions.md D-052 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | jt.cov60 |  | 78.4 | 1 |  | Mixed-fit calibration erratum disclosed in reports/joint_study.md |  | joint-d052-operational ¦ candidate 80 % range coverage, 60 min
C-53-01 | D-053 S 60-min personalised MAE 22.00 | README.md (Results at a glance); README §7; Deck s11 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | cs.S.mae60 |  | 22 | 2 |  |  |  | constrained-d053 ¦ S 60-min personalised MAE
C-53-02 | D-053 S 30-min MAE 13.89 | README.md (Results at a glance); Deck s11 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | cs.S.mae30 |  | 13.89 | 2
C-53-03 | D-053 persistence 15.34 / 24.47 (30 min) | README.md (Results at a glance); Deck s11 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | cs.persistence.mae30 |  | 15.34 | 2
C-53-04 | D-053 persistence 15.34 / 24.47 (60 min) | README.md (Results at a glance); Deck s11 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | cs.persistence.mae60 |  | 24.47 | 2
C-53-05 | D-053 R0 refitted in the same folds 23.20 | README.md (Results at a glance); narration metric card | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | cs.R0.mae60 |  | 23.2 | 2 |  |  |  | constrained-d053 ¦ R0 60-min personalised MAE (refitted in the same folds)
C-53-06 | D-053 S − R0 −1.19 at 60 min | README.md (Results at a glance); narration scene 11 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | cs.p60.mean_diff |  | -1.19 | 2 |  |  |  | constrained-d053 ¦ S − R0 at 60 min (participant bootstrap)
C-53-07 | D-053 S − R0 CI lower −2.09 | README.md (Results at a glance) | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | cs.p60.lo |  | -2.09 | 2
C-53-08 | D-053 S − R0 CI upper −0.39 | README.md (Results at a glance) | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | cs.p60.hi |  | -0.39 | 2
C-53-09 | D-053 lower for 11 of 14 people at 60 min | README §7; narration scene 11 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | cs.p60.improved |  | 11 | 0 |  |  |  | constrained-d053 ¦ people better at 60 min
C-53-10 | D-053 30-min change −0.50 | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | cs.pg.worsening_30 |  | -0.5 | 2 |  |  |  | constrained-d053 ¦ 30-min change
C-53-11 | D-053 error above 180 rose 1.07 against a +1.0 limit | README.md (Results at a glance); narration scene 11, 16 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | negative-result | cs.pg.high_worsening_60 |  | 1.07 | 2 |  | Unrounded +1.0658 > +1.0: gate fails at any display rounding |  | constrained-d053 ¦ high-target change at 60 min
C-53-12 | D-053 primary gate not met; not adopted | README.md (Results at a glance); docs/decisions.md D-053; narration scene 11 | Gate | negative-result | Gates: D-053; cs.pg.passed |  | FAILED |  |  |  |  | constrained-d053 ¦ primary gate
C-53-13 | D-053 false alarms per observed day S 1.15 | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | cs.w.S.fa_day |  | 1.15 | 2 |  |  |  | constrained-d053 ¦ false alarms per observed day, S
C-53-14 | D-053 false alarms per observed day R0-P 1.20 | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | cs.w.R0P.fa_day |  | 1.2 | 2 |  |  |  | constrained-d053 ¦ false alarms per observed day, R0-P
C-53-15 | D-053 onsets warned ≥ 30 min ahead, S 21.9 % | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | cs.w.S.share30 |  | 21.9 | 1 |  |  |  | constrained-d053 ¦ onsets warned ≥ 30 min ahead, S
C-53-16 | D-053 onsets warned ≥ 30 min ahead, R0-P 16.4 % | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | cs.w.R0P.share30 |  | 16.4 | 1 |  |  |  | constrained-d053 ¦ onsets warned ≥ 30 min ahead, R0-P
C-53-17 | D-053 operational gate passed | README §7 | Gate | supported | Gates: D-053 operational |  | PASSED |  |  |  |  | constrained-d053 ¦ operational gate
C-53-18 | D-053 range coverage 79.4 % at 30 min | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | cs.cov30 |  | 79.4 | 1 |  |  |  | constrained-d053 ¦ range coverage 30 min
C-53-19 | D-053 range coverage 77.3 % at 60 min | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | cs.cov60 |  | 77.3 | 1 |  |  |  | constrained-d053 ¦ range coverage 60 min
C-54-01 | D-054 foundation 60-min MAE 22.92 | README.md (Results at a glance); Deck s11; Deck s12; narration scene 16 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) — all origins | supported | fd.FM60 |  | 22.92 | 2 |  | Zero-shot, CGM only, not fine-tuned, not served |  | foundation-d054 ¦ foundation 60-min MAE
C-54-02 | D-054 foundation 30-min MAE 14.55 | README.md (Results at a glance); Deck s11 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | fd.FM30 |  | 14.55 | 2
C-54-03 | D-054 persistence 24.47 on identical origins | README.md (Results at a glance); Deck s11; narration scene 16 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | fd.P60 |  | 24.47 | 2 |  |  |  | foundation-d054 ¦ persistence 60-min MAE (identical origins)
C-54-04 | D-054 foundation − persistence −1.54 [−1.81, −1.30] | README §7; Deck s12 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | fd.fp.mean |  | -1.54 | 2 |  |  |  | foundation-d054 ¦ foundation − persistence
C-54-05 | D-054 all 14 people better than persistence | README §7; Deck s12; narration scene 11, 16 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | fd.fp.people_better |  | 14 | 0 |  |  |  | foundation-d054 ¦ people better than persistence
C-54-06 | D-054 refitted serving recipe 22.50 at 60 min | README.md (Results at a glance); Deck s12; narration scene 16 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | fd.R060 |  | 22.5 | 2 |  |  |  | foundation-d054 ¦ refitted serving recipe 60-min MAE (saved D-052 R0)
C-54-07 | D-054 foundation − refitted recipe +0.42 [−0.70, +1.47]: not better | README §7; Deck s12; narration scene 11 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | negative-result | fd.fr.mean |  | 0.42 | 2 |  |  |  | foundation-d054 ¦ foundation − refitted recipe
C-54-08 | D-054 F1 residual − same-code refit −0.46 [−0.72, −0.21] (descriptive) | README §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | fd.f1.mean |  | -0.46 | 2 |  | Descriptive secondary evidence; below the −0.5 bar used elsewhere; not promoted |  | foundation-d054 ¦ F1 residual − same-code refit (descriptive)
C-PR-01 | App replay reproduced all 15,440 saved test forecasts | README.md (Results at a glance); README Safety guardrails; Deck s10; narration metric card | Software test (saved parity run) | tested | parity.compared |  | 15440 | 0 |  | 15,440 = forecasts (7,720 issued origins × 2 horizons), not origins or people |  | parity ¦ forecasts compared (saved parity run, outcome pass for the serving model)
C-PR-02 | Parity tolerance 0.05 mg/dL | README.md (Results at a glance); Deck s10 | Software test | tested | parity.tolerance_mg_dl |  | 0.05 | 2 |  |  |  | parity ¦ tolerance (mg/dL)
C-SY-01 | Synthetic demonstration model is fitted on fictional data, never evidence | README (Public demonstration) | Synthetic mode | supported | reg.synthetic_fitted; registry | — | — |  | — |  |  | synthetic ¦ synthetic-demo.lightgbm.dyn-v2.v1 fitted on fictional data (registry synthetic_fitted)
C-FC-05 | Serving vs persistence −1.88 mg/dL at 30 min | README.md §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.cmp.sp.d30 |  | -1.88 | 2 |  | Candidate − control; participant-cluster bootstrap
C-FC-06 | Serving vs persistence 30-min CI [−2.53, −1.32] | README.md §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.cmp.sp.hi30; fus.cmp.sp.lo30 |  | -1.32 | 2
C-FC-07 | Serving vs persistence −2.96 [−4.34, −1.77] at 60 min | README.md §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.cmp.sp.d60 |  | -2.96 | 2
C-FC-08 | Serving better than persistence for 14/14 participants at 30 min | README.md §7; Deck s13 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.better30 |  | 14 | 0
C-FC-09 | Serving better than persistence for 13/14 participants at 60 min | README.md §7; Deck s13 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.better60 |  | 13 | 0
C-FC-10 | Patient adaptation −0.18 [−0.35, −0.01] at 30 min | README.md §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.cmp.adapt.d30 |  | -0.18 | 2 |  | Small
C-FC-11 | Patient adaptation −0.33 [−0.68, +0.02] at 60 min: inconclusive | README.md §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | negative-result | fus.cmp.adapt.d60 |  | -0.33 | 2 |  | CI crosses zero
C-FC-12 | LightGBM minus CatBoost +0.46 [+0.05, +0.92] at 60 min (population) | README.md §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.cmp.lgbcb.d60 |  | 0.46 | 2 |  | CatBoost better at 60 min; tie-break chose LightGBM
C-FC-13 | Profile vs no profile within the same 14 people −0.25 at 30 min | README.md §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.cmp.prof.d30 |  | -0.25 | 2 |  | May be identity, not physiology
C-FC-14 | Profile vs no profile within the same 14 people −0.61 at 60 min | README.md §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.cmp.prof.d60 |  | -0.61 | 2
C-FC-15 | CatBoost + profile + adaptation 12.93 / 19.80 (tied on validation, not selected) — 30 min | README.md §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.cbpa.mae30 |  | 12.93 | 2
C-FC-16 | CatBoost + profile + adaptation 19.80 at 60 min | README.md §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.cbpa.mae60 |  | 19.8 | 2
C-FC-17 | Previous demo model (v2) 13.15 / 20.81 — 30 min | README.md §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.v2a.mae30 |  | 13.15 | 2
C-FC-18 | Previous demo model (v2) 20.81 at 60 min | README.md §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.v2a.mae60 |  | 20.81 | 2
C-FC-19 | 7,730 forecast origins in the final test segment | README.md §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.seg.test.origins |  | 7730 | 0 |  | Origins ≠ issued origins (7,720) ≠ matched targets
C-FC-20 | ≈ 644 h of final-test recordings | README.md §7; narration metric card | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.seg.test.covered_hours |  | 644 | 0
C-WN-07 | Precision 0.60 | README.md §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.ev.precision |  | 0.6 | 2
C-WN-08 | 2.28 false alarms per eligible day | README.md §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.ev.false_alarm_episodes; fus.ev.eligible_patient_days |  | 2.28 | 2 |  | Eligible patient-days differ from observed patient-days
C-WN-09 | 48.4 % of all onsets warned at least 30 min ahead | README.md (Results at a glance); reports/warning_utility.md; Deck s13; narration scene 6 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | wu.ge30; wu.onsets |  | 48.4 | 1 |  | Denominator = all 64 eligible onsets (missed onsets count against it)
C-WN-10 | Warned ≥ 15 min ahead 67.2 % of all onsets | reports/warning_utility.md | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | wu.ge15; wu.onsets |  | 67.2 | 1
C-WN-11 | Warned ≥ 60 min ahead 20.3 % of all onsets | reports/warning_utility.md | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | wu.ge60; wu.onsets |  | 20.3 | 1
C-WN-12 | Qualified lead IQR 10–55 min (detected onsets) | reports/warning_utility.md | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | wu.lead_q75; wu.lead_q25 |  | 55 | 0
C-WN-13 | 40 of 101 warning episodes were false alarms | README (GIF caption) | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.ev.warning_episodes |  | 101 | 0
C-WN-14 | Persistence-based warning: 63/64 detected with 70 false alarms (2.61 per observed day) | README.md §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.pev.false_alarms_per_observed_patient_day |  | 2.61 | 2
C-RG-05 | Ranges covered only 60–82 % (30 min) per participant (min) | README.md §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.iv.per_patient_min30 |  | 60 | 0
C-RG-06 | Ranges per participant 30-min max 82 % | README.md §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.iv.per_patient_max30 |  | 82 | 0
C-RG-07 | Ranges per participant 60-min 51–88 % (min) | README.md §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.iv.per_patient_min60 |  | 51 | 0
C-RG-08 | Ranges per participant 60-min max 88 % | README.md §7 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.iv.per_patient_max60 |  | 88 | 0
C-RG-09 | While glucose rose (> 1 mg/dL/min) served ranges covered 51.6 % at 60 min on test | docs/decisions.md D-050 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | rg.inc.rising60 |  | 51.6 | 1 |  | Final-test trend context (read once in D-050)
C-TR-01 | Validation: 60-min error 35.3 mg/dL while glucose rose > 2 mg/dL/min | README §5; narration scene 15 | Serving model, validation segment | supported | tr.rising_fast.mae60 |  | 35.3 | 1 |  | Validation segment, not the test
C-TR-02 | Validation: 60-min error 19.5 when steady | README §5; narration scene 15 | Serving model, validation segment | supported | tr.steady.mae60 |  | 19.5 | 1
C-TR-03 | Peaks above 180 under-forecast by 16.9 mg/dL at 60 min (validation) | README §5; narration scene 15 | Serving model, validation segment | supported | tr.bias_high60 |  | 16.9 | 1
C-EX-06 | External 30-min difference −0.99 [−1.21, −0.77], better for 80 of 98 | README.md §7 | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | ext.cmp.mean_diff30 |  | -0.99 | 2
C-EX-07 | External better for 80 of 98 at 30 min | README.md §7 | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | ext.cmp.improved30 |  | 80 | 0
C-EX-08 | External 60-min difference −0.34 [−0.76, +0.10]: inconclusive | README.md §7 | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | negative-result | ext.cmp.mean_diff60 |  | -0.34 | 2
C-EX-09 | External ranges covered 87.1 % / 86.1 % (30 min) | README.md §7 | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | ext.cov30 |  | 87.1 | 1
C-EX-10 | External ranges 86.1 % at 60 min | README.md §7 | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | ext.cov60 |  | 86.1 | 1
C-NS-06 | v2 recipe refit per fold 14.52 / 24.25 (30 min) | README.md §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | nest.v2.mae30 |  | 14.52 | 2
C-NS-07 | v2 recipe refit per fold 24.25 at 60 min | README.md §7 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | nest.v2.mae60 |  | 24.25 | 2
C-2H-12 | Two-hour selected minus persistence −5.25 [−6.44, −4.05] | docs/evidence_checklist_2026-10-20.md | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | 2h.cmp.d |  | -5.25 | 2
C-GD-01 | No forecast is issued with less than 60 min of history (abstains) | README Safety guardrails; Deck s10 | Serving software | tested | tests.test_replay_api::test_insufficient_history_abstains_with_nulls | — | — |  | —
C-GD-02 | After 15 min without a reading the last forecast is shown marked Stale | README Safety guardrails; Deck s10 | Serving software | tested | cfg.stale; e2e: Participant 12: stale glucose before the gap closes, then abstention |  | 15 | 0
C-GD-03 | Correction needs ≥ 6 matured outcomes, capped at ±25 mg/dL, resets after a 3-hour gap | README Safety guardrails; Deck s10 | Serving software | tested | cfg.reset_gap; tests.test_replay_api::test_personalization_reports_its_bound_and_reset_gap |  | 3 | 0
C-GD-04 | Every warning is labelled Research warning and stays Outcome pending until it can be judged | README Safety guardrails; Deck s10 | Serving software | tested | e2e: Participant 01: a warning is pending until its outcome is knowable | — | — |  | —
C-GD-05 | Opened tests write a lock file; re-running needs a written reason | README Safety guardrails; Deck s10 | Research software | tested | tests.test_research_pipeline::test_editing_a_frozen_protocol_is_refused | — | — |  | —
C-GD-06 | Event = new episode above 180 mg/dL sustained ≥ 15 minutes | README §4 | Event definition | implemented | cfg.threshold; cfg.confirm; tests.test_events::test_exact_cutoff_is_not_high |  | 180 | 0
C-AU-01 | CI green on both branches | context.md (handoff) | CI | author-reported |  | — | — |  | — | No saved CI log in the repository | Link the GitHub Actions run after the repository is public
C-AU-02 | Fresh-clone checks on 23 Oct: pytest pass, vitest 141/141, e2e 16 passed / 13 skipped (synthetic) | Team handoff (23 Oct) | Fresh clone, synthetic data | author-reported |  | — | — |  | — | Logs not saved; not independently verified | Save a ledger for the fresh-clone run, or cite only the f073a10 ledger
C-VR-01 | pytest 221 passed, 2 skipped at f073a10 | context.md; reports/verification/final_2026-10-22_145326Z/LEDGER.md | Local verification at f073a10 (not later revisions) | tested | ver.pytest.passed; ver.pytest.skipped |  | 221 | 0
C-VR-02 | vitest 141 passed at f073a10 | context.md; LEDGER.md | Local verification at f073a10 | tested | ver.vitest.passed |  | 141 | 0
C-VR-03 | e2e 29 normal/recorded at f073a10 | context.md; LEDGER.md | Local verification at f073a10 | tested | ver.e2e.passed |  | 29 | 0
C-VR-04 | e2e 4 public at f073a10 | LEDGER.md | Local verification at f073a10 | tested | ver.e2e_public.passed |  | 4 | 0
C-VR-05 | e2e 3 summary/export at f073a10 | LEDGER.md | Local verification at f073a10 | tested | ver.e2e_summary.passed |  | 3 | 0
C-VR-06 | Parity re-run in the final ledger: 15,440 compared, 0 over tolerance | LEDGER.md | Local verification at f073a10 | tested | parity_final.compared; parity_final.over_tolerance |  | 15440 | 0
C-S1-01 | GlucoTwin — Team Kryptonite, VESIT; Avinash Gehi (team leader) · Hardik Hinduja | Slide 1; Deck s15 | Team facts | supported | submission (configs/submission.yaml team.*) | — | — |  | —
C-S1-02 | Digital Twin Challenge 2026 · Happiest Health | Slide 1; README | Challenge identity | author-reported |  | — | — |  | — | Challenge page not archived in the repository | Optional: cite the Unstop challenge URL
C-S1-03 | Student research prototype — not a medical device | Slide 1; README | Scope statement | author-reported |  | — | — |  | — | A declaration of intended use; no regulatory submission exists or is claimed | See limitation
C-S2-01 | HbA1c summarises ~3 months of glucose; it cannot show short high-glucose excursions (NIDDK) | Slide 2; README §4 | External fact | pending |  | — | — |  | — | No NIDDK URL is cited anywhere in the repository | Pending primary-source check: locate the NIDDK page and record its URL and access date
C-S2-02 | WHO identifies continuity of care and diabetes monitoring as ongoing challenges in India | Slide 2 | External fact | pending |  | — | — |  | — | No WHO document is cited anywhere in the repository | Pending primary-source check: locate the WHO source and record its URL and access date
C-S2-03 | Continuous glucose monitors produce ~288 readings a day | Slide 2; README §4–5 | Arithmetic | supported-arithmetic | const.min_per_day; const.cadence |  | 288 | 0 |  | Assumes a 5-minute sensor cadence (Dexcom G6 Pro)
C-S2-04 | Question for the clinician (framing) | Slide 2 | Framing | proposed/illustrative |  | — | — |  | —
C-S2-05 | Target: forecast glucose 30 and 60 minutes ahead and warn of a NEW sustained episode > 180 mg/dL (≥ 15 min), Type 2 | Slide 2; README §4 | Task definition | implemented | cfg.confirm; cfg.threshold; fus.participants |  | 15 | 0
C-S3-01 | Hypothetical Mumbai diabetes educator illustration | Slide 3; README §5 | Illustration | proposed/illustrative |  | — | — |  | — | No clinic, partner, pilot or Indian participants
C-S3-02 | Proposed users and workflow | Slide 3 | Proposal | proposed/illustrative |  | — | — |  | —
C-S3-03 | Never: no insulin advice, triage, live device claims or LLM predictions | Slide 3; README; Deck s10 | Scope boundary | implemented |  | — | — |  | — | Absence of features verified by code inspection, not by a dedicated test
C-S3-04 | Potential fit with diagnostics + outpatient services is a proposal — no partnership or pilot is claimed | Slide 3 | Proposal | proposed/illustrative |  | — | — |  | —
C-S4-01 | Versioned patient state | Slide 4 | Serving software | tested | tests.test_replay_api::test_review_snapshot_is_one_session_revision_and_clock | — | — |  | —
C-S4-02 | Replay clock releases only what was available | Slide 4; Deck s10 | Serving software | tested | tests.test_replay_api::test_event_polling_has_no_duplicates_or_gaps_and_respects_availability; tests.test_leakage::test_late_logged_meal_is_hidden_until_available | — | — |  | —
C-S4-03 | Population forecaster + patient correction learned from matured errors | Slide 4 | Serving software | tested | tests.test_leakage::test_sixty_minute_residual_matures_only_when_target_is_available; tests.test_leakage::test_corrections_never_stack_on_corrected_predictions | — | — |  | —
C-S4-04 | A statistical research twin, not a physiological simulator | Slide 4 | Definition | implemented |  | — | — |  | — | Describes the method class (no physiological model in code)
C-S5-01 | CGMacros 1.0.0 (PhysioNet, CC BY-NC-SA 4.0) | Slide 5; docs/data_card.md; Deck s15 | Dataset licence | supported | src.cgm.license; src.cgm.doi | — | — |  | —
C-S5-02 | CGMacros: 45 adults | Slide 5 | Dataset | supported | inv.cgm |  | 45 | 0
C-S5-03 | ~10 days, Dexcom G6 Pro + Libre Pro (blinded), Fitbit, meals, screening labs | Slide 5; docs/data_card.md | Dataset description | supported | data_card | — | — |  | — | From the publisher description recorded in the data card; not re-checked against PhysioNet in this build
C-S5-04 | Primary cohort: 14 participants in the Type 2 group (HbA1c > 6.4 %) | Slide 5; README | Cohort | supported | fus.participants; decisions (D-006) |  | 14 | 0
C-S5-05 | 31 others replayable but excluded | Slide 5; docs/data_card.md | Cohort | supported | inv.cgm; fus.participants; fus.excluded |  | 31 | 0
C-S5-06 | Published 1-minute CGM columns are linearly interpolated; native 5-minute Dexcom readings recovered | Slide 5; Deck s13 | Data handling | tested | data_audit; tests.test_leakage::test_one_minute_interpolated_rows_do_not_inflate_observations | — | — |  | —
C-S5-07 | Undocumented columns reported; dates shifted; timezone unknown → synthetic clock, no time-of-day features | Slide 5 | Data handling | supported | data_audit (Undocumented columns; Time) | — | — |  | — | Data card says the synthetic timeline preserves source time of day; 'no time-of-day features' is a model-input statement checked against the feature list, not a test
C-S5-08 | Synthetic fictional participants (labelled synthetic) | Slide 5; README | Synthetic mode | tested | tests.test_public_demo::test_synthetic_only; reg.synthetic_fitted | — | — |  | —
C-D6-01 | A forecast at time t uses only records observed and available by t | Deck s6 | Serving software | tested | tests.test_leakage::test_future_values_cannot_change_earlier_features; tests.test_data_expansion::test_a_reading_released_late_is_invisible_before_its_availability | — | — |  | —
C-D6-02 | Meals left out of the main model (logging time unknown); tested separately | Deck s6 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.sens.meal60; fus.sens.meal60; data_card |  | — |  | see limitation | Optimistic meal-logging sensitivity, not eligible for serving
C-D6-03 | Protocol frozen before fitting; choices made on validation | Deck s6 | Protocol | supported | Studies sheet (protocol first commit vs run); fus.selection_rule | — | — |  | — | The fusion protocol's first commit is 25 s after the run's created_at and its constraint was declared after the test was inspected (disclosed)
C-D6-04 | The fusion re-run (gt-primary-2026-10-13-fusion) was declared after gt-primary-2026-10-11-v2's test had been inspected | Deck s6; Deck s9 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | decisions (D-038/D-039); fusion /correction_note | — | — |  | —
C-D7-01 | Synthetic patients: 5 fictional profiles (CC0) | Deck s7 | Synthetic mode | implemented | ml/glucotwin/data/synthetic.py (scenarios SYN-A…SYN-E, LICENSE CC0-1.0); docs/data_card.md | — | — |  | —
C-D7-02 | External cohort ShanghaiT2DM, 98 people, 15-min CGM, research only | Deck s7 | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | inv.sh_eval |  | 98 | 0
C-D7-03 | Replay releases a reading once available_at ≤ t; seek / reset → new revision | Deck s7 | Serving software | tested | tests.test_replay_api::test_event_polling_has_no_duplicates_or_gaps_and_respects_availability; tests.test_replay_api::test_seek_backward_removes_future_state_and_invalidates_cursors | — | — |  | —
C-D7-04 | Causal features, ≤ 180 min history | Deck s7 | Serving software | implemented | cfg.lookback |  | 180 | 0
C-D7-05 | Fusion LightGBM: 8 profile + 38 dynamic features | Deck s7; Deck s8 | Serving model | author-reported |  | — | — |  | — | Counts stated in docs/model_card.md and README; not re-derived from the model artifact by this builder | Optional: read the feature list from the registered model bundle
C-D7-06 | Split 60/10/10/20 with purged boundaries | Deck s7 | Protocol | tested | tests.test_leakage::test_split_purging_keeps_label_support_inside_segments; configs/experiment.yaml | — | — |  | —
C-D7-07 | CPU only, local, no cloud, no LLM; not a medical device | Deck s7 | Scope boundary | implemented |  | — | — |  | — | Absence verified by code inspection only
C-D8-01 | Baselines persistence, linear trend, ridge; learners LightGBM and CatBoost, same search budget | Deck s8 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.selection_rule; selection candidates (Sources sel.*) | — | — |  | —
C-D8-02 | Patient correction: a bounded mean of this person's observed errors; model weights never refit | Deck s8 | Serving software | tested | tests.test_leakage::test_corrections_never_stack_on_corrected_predictions; tests.test_replay_api::test_personalization_reports_its_bound_and_reset_gap | — | — |  | —
C-D8-03 | 80 % prediction ranges (split-conformal), calibrated for the exact serving pipeline | Deck s8 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | const.target_cov; fus.iv.q30; fus.iv.q60 |  | 80 | 0
C-D8-04 | Warning score: highest forecast in the next 60 min minus 180 mg/dL — not a probability | Deck s8 | Serving software | implemented | cfg.threshold |  | 180 | 0
C-D9-10 | Rule: lowest mean per-participant validation MAE over both horizons; within 0.5 mg/dL → simpler model | Deck s9 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.selection_rule | — | — |  | —
C-D9-11 | Only dyn-prof-v2 eligible (fusion constraint added 2026-10-13, after the test segment had been inspected) | Deck s9 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.selection_rule; decisions (D-038) | — | — |  | —
C-D9-12 | Selected: lightgbm:dyn-prof-v2 with patient correction | Deck s9 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | sel.selected; reg.serving | — | — |  | —
C-D9-13 | Without the constraint the rule picks lightgbm:dyn-v2 (gt-primary-2026-10-11-v2) | Deck s9 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | sel.v2_selected | — | — |  | —
C-D12-01 | amazon/chronos-bolt-tiny, 8.7 M parameters, apache-2.0, zero-shot on CGM only | Deck s12 | Research baseline (D-054) | supported | fd.params; fd.checkpoint; fd.license |  | 8.7 | 1
C-D12-02 | Foundation vs persistence −0.79 [−1.01, −0.59] mg/dL at 30 min | Deck s12 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | fd.fp30.mean |  | -0.79 | 2
C-D12-03 | 30-min CI lower −1.01 | Deck s12 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | fd.fp30.lo |  | -1.01 | 2
C-D12-04 | 30-min CI upper −0.59 | Deck s12 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | fd.fp30.hi |  | -0.59 | 2
C-D12-05 | Foundation vs persistence 60-min CI lower −1.81 | Deck s12 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | fd.fp.lo |  | -1.81 | 2
C-D12-06 | Foundation vs persistence 60-min CI upper −1.30 | Deck s12 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | fd.fp.hi |  | -1.3 | 2
C-D12-07 | Foundation vs refitted recipe CI [−0.70, +1.47] (lower) | Deck s12 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | fd.fr.lo |  | -0.7 | 2
C-D12-08 | Foundation vs refitted recipe CI upper +1.47 | Deck s12 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | supported | fd.fr.hi |  | 1.47 | 2
C-D12-09 | Foundation worse above 180 mg/dL than the refitted recipe | Deck s12 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | negative-result | fd.high.FM; fd.high.R0; fd.high.FM; fd.high.R0 |  | — |  | see limitation | Positive difference = foundation worse
C-D12-10 | Its pretraining data may include public CGM recordings | Deck s12 | Disclosure | author-reported |  | — | — |  | — | A caution, not a measured fact | See limitation
C-D13-01 | Built: replay API, immutable forecasts, warnings with outcomes, clinician workspace | Deck s13 | Serving software | tested | tests.test_replay_api::test_forward_replay_equals_reset_and_seek; e2e: Participant 01: a warning is pending until its outcome is knowable | — | — |  | —
C-D13-02 | Reproducible: synthetic mode needs no download; results regenerate from versioned configs | Deck s13 | Software | tested | tests.test_replay_api::test_synthetic_demo_runs_without_dataset | — | — |  | —
C-D14-01 | 14 participants from one US site; no Indian participants | Deck s14 | Limitation | supported | fus.participants; data_card |  | 14 | 0
C-D14-02 | Blinded CGMs replayed with assumed zero latency; medication and meal timing unknown | Deck s14 | Limitation | supported | data_card; docs/model_card.md | — | — |  | —
C-D14-03 | Ranges under-cover (74.2 % / 74.0 % vs 80 % target) | Deck s14; Deck s11 | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.iv.coverage30; const.target_cov; fus.iv.coverage60 |  | — |  | see limitation | Negative = under-coverage
C-D14-04 | The warning threshold ignores review workload | Deck s14 | Limitation | author-reported |  | — | — |  | — | Threshold chosen by event F1 on validation (implemented); the workload statement is a design observation | See limitation
C-D14-05 | Retrospective replay only — no clinical, workflow or safety study | Deck s14 | Limitation | supported | README; docs/model_card.md | — | — |  | —
C-D14-06 | Next: a prospective, consented Indian cohort and a full-fusion evaluation on new people | Deck s14 | Proposal | proposed/illustrative |  | — | — |  | —
C-D15-01 | Code: MIT license (code only); quick start make setup · make synthetic · make dev | Deck s15 | Repository | supported | submission (license.software MIT); LICENSE; Makefile | — | — |  | —
C-D15-02 | Public GitHub link and demo video added after publication/upload | Deck s15 | Submission status | pending |  | — | — |  | — |  | Fill links.repository and links.video in configs/submission.yaml after the team publishes and uploads
C-D15-03 | ShanghaiT2DM (figshare, doi:10.6084/m9.figshare.21600933.v5, CC BY 4.0) | Deck s15 | Dataset licence | supported | src.sh.doi; src.sh.license | — | — |  | —
C-NR-01 | Per-person 60-min range coverage 51.4 %–87.6 % (min) | Narration metric card | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.iv.per_patient_min60 |  | 51.4 | 1
C-NR-02 | Per-person 60-min range coverage max 87.6 % | Narration metric card | Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | supported | fus.iv.per_patient_max60 |  | 87.6 | 1
C-NR-03 | Range held 46 % while rising > 2 mg/dL/min at 60 min (validation) | Narration scene 15; metric card | Serving model, validation segment | supported | tr.rising_fast.cov60 |  | 46 | 0
C-NR-04 | Joint selection lowered 60-min error but made peaks above 180 worse and raised false alarms: failed its gates | Narration scene 11 | Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) | negative-result | jt.rel.fa_per_day_delta; Gates: D-052; jt.gate.high_worsening_60 |  | — |  | see limitation
C-NR-05 | D-053 passed the warning and range checks | Narration scene 11, 16 | Gate | supported | Gates: D-053 operational; cs.rg.passed | — | — |  | —
C-NR-06 | Blending, retraining, tuning and joint selection each failed their declared bars | Narration scene 16 | Gate | negative-result | Gates: D-048, D-049, D-051, D-052 | — | — |  | —
C-NR-07 | Shanghai: a separate CGM-only model beat persistence at 30 min and at two hours | Narration scene 16 | External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) | supported | 2hx.cmp.hi; ext.cmp.hi30 |  | — |  | see limitation | 120-min CI upper < 0; 30-min D-032 CI upper < 0
C-D9-0130 | Slide 9 table: lightgbm:dyn-prof-v2 val. MAE 30 | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.lightgbm_dyn-prof-v2.selection_mae_30 |  | 13.47 | 2
C-D9-0160 | Slide 9 table: lightgbm:dyn-prof-v2 val. MAE 60 | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.lightgbm_dyn-prof-v2.selection_mae_60 |  | 21.26 | 2
C-D9-01sc | Slide 9 table: lightgbm:dyn-prof-v2 score | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.lightgbm_dyn-prof-v2.selection_score |  | 17.36 | 2
C-D9-0230 | Slide 9 table: catboost:dyn-prof-v2 val. MAE 30 | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.catboost_dyn-prof-v2.selection_mae_30 |  | 13.56 | 2
C-D9-0260 | Slide 9 table: catboost:dyn-prof-v2 val. MAE 60 | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.catboost_dyn-prof-v2.selection_mae_60 |  | 21.31 | 2
C-D9-02sc | Slide 9 table: catboost:dyn-prof-v2 score | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.catboost_dyn-prof-v2.selection_score |  | 17.44 | 2
C-D9-0330 | Slide 9 table: lightgbm:dyn-v2 val. MAE 30 | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.lightgbm_dyn-v2.selection_mae_30 |  | 13.76 | 2
C-D9-0360 | Slide 9 table: lightgbm:dyn-v2 val. MAE 60 | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.lightgbm_dyn-v2.selection_mae_60 |  | 21.73 | 2
C-D9-03sc | Slide 9 table: lightgbm:dyn-v2 score | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.lightgbm_dyn-v2.selection_score |  | 17.75 | 2
C-D9-0430 | Slide 9 table: catboost:dyn-v2 val. MAE 30 | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.catboost_dyn-v2.selection_mae_30 |  | 13.88 | 2
C-D9-0460 | Slide 9 table: catboost:dyn-v2 val. MAE 60 | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.catboost_dyn-v2.selection_mae_60 |  | 21.66 | 2
C-D9-04sc | Slide 9 table: catboost:dyn-v2 score | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.catboost_dyn-v2.selection_score |  | 17.77 | 2
C-D9-0530 | Slide 9 table: ridge:dyn-v2 val. MAE 30 | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.ridge_dyn-v2.selection_mae_30 |  | 14.13 | 2
C-D9-0560 | Slide 9 table: ridge:dyn-v2 val. MAE 60 | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.ridge_dyn-v2.selection_mae_60 |  | 22.44 | 2
C-D9-05sc | Slide 9 table: ridge:dyn-v2 score | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.ridge_dyn-v2.selection_score |  | 18.28 | 2
C-D9-0630 | Slide 9 table: ridge:dyn-prof-v2 val. MAE 30 | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.ridge_dyn-prof-v2.selection_mae_30 |  | 14.21 | 2
C-D9-0660 | Slide 9 table: ridge:dyn-prof-v2 val. MAE 60 | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.ridge_dyn-prof-v2.selection_mae_60 |  | 22.61 | 2
C-D9-06sc | Slide 9 table: ridge:dyn-prof-v2 score | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.ridge_dyn-prof-v2.selection_score |  | 18.41 | 2
C-D9-0730 | Slide 9 table: linear_trend:cgm-only-v1 val. MAE 30 | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.linear_trend_cgm-only-v1.selection_mae_30 |  | 15.66 | 2
C-D9-0760 | Slide 9 table: linear_trend:cgm-only-v1 val. MAE 60 | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.linear_trend_cgm-only-v1.selection_mae_60 |  | 26.33 | 2
C-D9-07sc | Slide 9 table: linear_trend:cgm-only-v1 score | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.linear_trend_cgm-only-v1.selection_score |  | 20.99 | 2
C-D9-0830 | Slide 9 table: persistence:cgm-only-v1 val. MAE 30 | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.persistence_cgm-only-v1.selection_mae_30 |  | 16.51 | 2
C-D9-0860 | Slide 9 table: persistence:cgm-only-v1 val. MAE 60 | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.persistence_cgm-only-v1.selection_mae_60 |  | 26.32 | 2
C-D9-08sc | Slide 9 table: persistence:cgm-only-v1 score | Deck s9 | Serving experiment, validation segment (selection) | supported | sel.persistence_cgm-only-v1.selection_score |  | 21.42 | 2

## Sources
Sources — every source value stored once, at full precision |  |  |  |  | ← About / navigation
Value column is the only place a source number is typed by the builder; every other sheet references it. 'unavailable' = the file is not on this installation (never 0).
Value ID | Description | Value | Unit | Source ID | Repository path | Locator (JSON pointer / field) | Source date | Source SHA-256 | Protocol / revision identity | Publication status
fus.pers.mae30 | Serving test (final test segment): Persistence (last reading), 30 min, mae (macro over participants) | 14.8366 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=persistence:cgm-only-v1&personalization=off&split=test&horizon=30]/mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.pers.rmse30 | Serving test: Persistence (last reading), 30 min, rmse | 22.084 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=persistence:cgm-only-v1&personalization=off&split=test&horizon=30]/rmse | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.pers.mae60 | Serving test (final test segment): Persistence (last reading), 60 min, mae (macro over participants) | 23.2254 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=persistence:cgm-only-v1&personalization=off&split=test&horizon=60]/mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.pers.rmse60 | Serving test: Persistence (last reading), 60 min, rmse | 33.5582 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=persistence:cgm-only-v1&personalization=off&split=test&horizon=60]/rmse | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.trend.mae30 | Serving test (final test segment): Linear trend, 30 min, mae (macro over participants) | 14.0315 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=linear_trend:cgm-only-v1&personalization=off&split=test&horizon=30]/mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.trend.rmse30 | Serving test: Linear trend, 30 min, rmse | 21.0348 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=linear_trend:cgm-only-v1&personalization=off&split=test&horizon=30]/rmse | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.trend.mae60 | Serving test (final test segment): Linear trend, 60 min, mae (macro over participants) | 23.2313 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=linear_trend:cgm-only-v1&personalization=off&split=test&horizon=60]/mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.trend.rmse60 | Serving test: Linear trend, 60 min, rmse | 33.9593 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=linear_trend:cgm-only-v1&personalization=off&split=test&horizon=60]/rmse | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.ridgep.mae30 | Serving test (final test segment): Ridge, dynamic + profile, 30 min, mae (macro over participants) | 14.6067 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=ridge:dyn-prof-v2&personalization=off&split=test&horizon=30]/mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.ridgep.rmse30 | Serving test: Ridge, dynamic + profile, 30 min, rmse | 21.2487 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=ridge:dyn-prof-v2&personalization=off&split=test&horizon=30]/rmse | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.ridgep.mae60 | Serving test (final test segment): Ridge, dynamic + profile, 60 min, mae (macro over participants) | 22.9644 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=ridge:dyn-prof-v2&personalization=off&split=test&horizon=60]/mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.ridgep.rmse60 | Serving test: Ridge, dynamic + profile, 60 min, rmse | 31.8311 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=ridge:dyn-prof-v2&personalization=off&split=test&horizon=60]/rmse | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.v2a.mae30 | Serving test (final test segment): LightGBM dynamic-only + correction (previous demo model, v2), 30 min, mae (macro over participants) | 13.1469 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=lightgbm:dyn-v2&personalization=adaptive&split=test&horizon=30]/mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.v2a.rmse30 | Serving test: LightGBM dynamic-only + correction (previous demo model, v2), 30 min, rmse | 19.2605 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=lightgbm:dyn-v2&personalization=adaptive&split=test&horizon=30]/rmse | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.v2a.mae60 | Serving test (final test segment): LightGBM dynamic-only + correction (previous demo model, v2), 60 min, mae (macro over participants) | 20.8074 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=lightgbm:dyn-v2&personalization=adaptive&split=test&horizon=60]/mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.v2a.rmse60 | Serving test: LightGBM dynamic-only + correction (previous demo model, v2), 60 min, rmse | 29.0689 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=lightgbm:dyn-v2&personalization=adaptive&split=test&horizon=60]/rmse | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.lgbp.mae30 | Serving test (final test segment): LightGBM + profile, population, 30 min, mae (macro over participants) | 13.1358 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=lightgbm:dyn-prof-v2&personalization=off&split=test&horizon=30]/mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.lgbp.rmse30 | Serving test: LightGBM + profile, population, 30 min, rmse | 18.8586 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=lightgbm:dyn-prof-v2&personalization=off&split=test&horizon=30]/rmse | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.lgbp.mae60 | Serving test (final test segment): LightGBM + profile, population, 60 min, mae (macro over participants) | 20.591 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=lightgbm:dyn-prof-v2&personalization=off&split=test&horizon=60]/mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.lgbp.rmse60 | Serving test: LightGBM + profile, population, 60 min, rmse | 28.136 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=lightgbm:dyn-prof-v2&personalization=off&split=test&horizon=60]/rmse | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.serv.mae30 | Serving test (final test segment): LightGBM + profile + correction (serving), 30 min, mae (macro over participants) | 12.9563 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=lightgbm:dyn-prof-v2&personalization=adaptive&split=test&horizon=30]/mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.serv.rmse30 | Serving test: LightGBM + profile + correction (serving), 30 min, rmse | 18.7907 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=lightgbm:dyn-prof-v2&personalization=adaptive&split=test&horizon=30]/rmse | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.serv.n_origins30 | Serving test: LightGBM + profile + correction (serving), 30 min, n_origins | 7699 | origins | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=lightgbm:dyn-prof-v2&personalization=adaptive&split=test&horizon=30]/n_origins | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.serv.n_patients30 | Serving test: LightGBM + profile + correction (serving), 30 min, n_patients | 14 | people | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=lightgbm:dyn-prof-v2&personalization=adaptive&split=test&horizon=30]/n_patients | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.serv.mae60 | Serving test (final test segment): LightGBM + profile + correction (serving), 60 min, mae (macro over participants) | 20.2643 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=lightgbm:dyn-prof-v2&personalization=adaptive&split=test&horizon=60]/mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.serv.rmse60 | Serving test: LightGBM + profile + correction (serving), 60 min, rmse | 28.0039 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=lightgbm:dyn-prof-v2&personalization=adaptive&split=test&horizon=60]/rmse | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.serv.n_origins60 | Serving test: LightGBM + profile + correction (serving), 60 min, n_origins | 7690 | origins | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=lightgbm:dyn-prof-v2&personalization=adaptive&split=test&horizon=60]/n_origins | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.serv.n_patients60 | Serving test: LightGBM + profile + correction (serving), 60 min, n_patients | 14 | people | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=lightgbm:dyn-prof-v2&personalization=adaptive&split=test&horizon=60]/n_patients | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cbpa.mae30 | Serving test (final test segment): CatBoost + profile + correction (not selected), 30 min, mae (macro over participants) | 12.9303 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=catboost:dyn-prof-v2&personalization=adaptive&split=test&horizon=30]/mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cbpa.rmse30 | Serving test: CatBoost + profile + correction (not selected), 30 min, rmse | 18.7301 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=catboost:dyn-prof-v2&personalization=adaptive&split=test&horizon=30]/rmse | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cbpa.mae60 | Serving test (final test segment): CatBoost + profile + correction (not selected), 60 min, mae (macro over participants) | 19.8028 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=catboost:dyn-prof-v2&personalization=adaptive&split=test&horizon=60]/mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cbpa.rmse60 | Serving test: CatBoost + profile + correction (not selected), 60 min, rmse | 27.7207 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /forecast_results[label=catboost:dyn-prof-v2&personalization=adaptive&split=test&horizon=60]/rmse | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.sp.d30 | Serving test paired difference, serving minus persistence, 30 min (A minus B; negative = A lower error) | -1.8803 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=serving minus persistence&horizon=30]/mean_diff_mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.sp.lo30 | 95 % CI lower (participant bootstrap) | -2.5333 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=serving minus persistence&horizon=30]/ci95[0] | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.sp.hi30 | 95 % CI upper (participant bootstrap) | -1.3168 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=serving minus persistence&horizon=30]/ci95[1] | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.sp.d60 | Serving test paired difference, serving minus persistence, 60 min (A minus B; negative = A lower error) | -2.9611 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=serving minus persistence&horizon=60]/mean_diff_mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.sp.lo60 | 95 % CI lower (participant bootstrap) | -4.339 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=serving minus persistence&horizon=60]/ci95[0] | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.sp.hi60 | 95 % CI upper (participant bootstrap) | -1.7676 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=serving minus persistence&horizon=60]/ci95[1] | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.adapt.d30 | Serving test paired difference, lightgbm:dyn-prof-v2: adaptive minus population, 30 min (A minus B; negative = A lower error) | -0.17947 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=lightgbm:dyn-prof-v2: adaptive minus population&horizon=30]/mean_diff_mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.adapt.lo30 | 95 % CI lower (participant bootstrap) | -0.35379 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=lightgbm:dyn-prof-v2: adaptive minus population&horizon=30]/ci95[0] | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.adapt.hi30 | 95 % CI upper (participant bootstrap) | -0.01492 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=lightgbm:dyn-prof-v2: adaptive minus population&horizon=30]/ci95[1] | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.adapt.d60 | Serving test paired difference, lightgbm:dyn-prof-v2: adaptive minus population, 60 min (A minus B; negative = A lower error) | -0.32663 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=lightgbm:dyn-prof-v2: adaptive minus population&horizon=60]/mean_diff_mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.adapt.lo60 | 95 % CI lower (participant bootstrap) | -0.67578 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=lightgbm:dyn-prof-v2: adaptive minus population&horizon=60]/ci95[0] | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.adapt.hi60 | 95 % CI upper (participant bootstrap) | 0.016871 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=lightgbm:dyn-prof-v2: adaptive minus population&horizon=60]/ci95[1] | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.lgbcb.d30 | Serving test paired difference, lightgbm minus catboost (dyn-prof-v2, population), 30 min (A minus B; negative = A lower error) | 0.022549 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=lightgbm minus catboost (dyn-prof-v2, population)&horizon=30]/mean_diff_mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.lgbcb.lo30 | 95 % CI lower (participant bootstrap) | -0.14807 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=lightgbm minus catboost (dyn-prof-v2, population)&horizon=30]/ci95[0] | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.lgbcb.hi30 | 95 % CI upper (participant bootstrap) | 0.19797 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=lightgbm minus catboost (dyn-prof-v2, population)&horizon=30]/ci95[1] | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.lgbcb.d60 | Serving test paired difference, lightgbm minus catboost (dyn-prof-v2, population), 60 min (A minus B; negative = A lower error) | 0.46315 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=lightgbm minus catboost (dyn-prof-v2, population)&horizon=60]/mean_diff_mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.lgbcb.lo60 | 95 % CI lower (participant bootstrap) | 0.049648 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=lightgbm minus catboost (dyn-prof-v2, population)&horizon=60]/ci95[0] | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.lgbcb.hi60 | 95 % CI upper (participant bootstrap) | 0.91859 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=lightgbm minus catboost (dyn-prof-v2, population)&horizon=60]/ci95[1] | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.prof.d30 | Serving test paired difference, lightgbm: profile minus dynamic-only (population), 30 min (A minus B; negative = A lower error) | -0.24756 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=lightgbm: profile minus dynamic-only (population)&horizon=30]/mean_diff_mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.prof.lo30 | 95 % CI lower (participant bootstrap) | -0.42992 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=lightgbm: profile minus dynamic-only (population)&horizon=30]/ci95[0] | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.prof.hi30 | 95 % CI upper (participant bootstrap) | -0.035878 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=lightgbm: profile minus dynamic-only (population)&horizon=30]/ci95[1] | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.prof.d60 | Serving test paired difference, lightgbm: profile minus dynamic-only (population), 60 min (A minus B; negative = A lower error) | -0.6081 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=lightgbm: profile minus dynamic-only (population)&horizon=60]/mean_diff_mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.prof.lo60 | 95 % CI lower (participant bootstrap) | -1.1484 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=lightgbm: profile minus dynamic-only (population)&horizon=60]/ci95[0] | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.cmp.prof.hi60 | 95 % CI upper (participant bootstrap) | -0.091152 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /comparisons[comparison=lightgbm: profile minus dynamic-only (population)&horizon=60]/ci95[1] | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.ev.true_onsets | Serving 60-min warning, final test: new sustained >180 onsets (eligible) | 64 | onsets | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /events[model=lightgbm:dyn-prof-v2 (adaptive)&split=test&horizon=60]/true_onsets | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.ev.detected_onsets | Serving 60-min warning, final test: onsets matched to a warning | 61 | onsets | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /events[model=lightgbm:dyn-prof-v2 (adaptive)&split=test&horizon=60]/detected_onsets | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.ev.warning_episodes | Serving 60-min warning, final test: warning episodes | 101 | episodes | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /events[model=lightgbm:dyn-prof-v2 (adaptive)&split=test&horizon=60]/warning_episodes | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.ev.false_alarm_episodes | Serving 60-min warning, final test: false-alarm episodes | 40 | episodes | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /events[model=lightgbm:dyn-prof-v2 (adaptive)&split=test&horizon=60]/false_alarm_episodes | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.ev.precision | Serving 60-min warning, final test: episode precision | 0.60396 | ratio | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /events[model=lightgbm:dyn-prof-v2 (adaptive)&split=test&horizon=60]/precision | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.ev.false_alarms_per_observed_patient_day | Serving 60-min warning, final test: saved false-alarm rate | 1.4922 | per observed patient-day | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /events[model=lightgbm:dyn-prof-v2 (adaptive)&split=test&horizon=60]/false_alarms_per_observed_patient_day | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.ev.observed_patient_days | Serving 60-min warning, final test: observed patient-day denominator (saved event metadata) | 26.8056 | patient-days | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /events[model=lightgbm:dyn-prof-v2 (adaptive)&split=test&horizon=60]/observed_patient_days | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.ev.eligible_patient_days | Serving 60-min warning, final test: eligible patient-day denominator (saved event metadata) | 17.5069 | patient-days | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /events[model=lightgbm:dyn-prof-v2 (adaptive)&split=test&horizon=60]/eligible_patient_days | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.ev.false_alarms_per_patient_day | Serving 60-min warning, final test: saved false alarms per eligible patient-day | 2.2848 | per eligible patient-day | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /events[model=lightgbm:dyn-prof-v2 (adaptive)&split=test&horizon=60]/false_alarms_per_patient_day | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.ev.qualified_lead_median | Serving 60-min warning, final test: median qualified lead among detected onsets | 30 | min | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /events[model=lightgbm:dyn-prof-v2 (adaptive)&split=test&horizon=60]/qualified_lead_median | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.ev.qualified_lead_n | Serving 60-min warning, final test: onsets with a qualified lead (denominator of the median) | 61 | onsets | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /events[model=lightgbm:dyn-prof-v2 (adaptive)&split=test&horizon=60]/qualified_lead_n | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.ev.threshold | Serving 60-min warning, final test: warning margin (score = max forecast − 180) | -10 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /events[model=lightgbm:dyn-prof-v2 (adaptive)&split=test&horizon=60]/threshold | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.ev.time_base | Serving 60-min warning, final test: time-base definition of the day denominators | approximate: number of forecast-issued origins x 5-minute native cadence | text | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /events[model=lightgbm:dyn-prof-v2 (adaptive)&split=test&horizon=60]/time_base | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.ev.iqr_lo | Qualified lead IQR lower (detected onsets) | 10 | min | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /events[model=lightgbm:dyn-prof-v2 (adaptive)&split=test&horizon=60]/qualified_lead_iqr[0] | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.ev.iqr_hi | Qualified lead IQR upper (detected onsets) | 55 | min | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /events[model=lightgbm:dyn-prof-v2 (adaptive)&split=test&horizon=60]/qualified_lead_iqr[1] | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.pev.detected_onsets | Persistence 60-min warning (fires at 150 mg/dL), final test: detected_onsets | 63 | count | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /events[model=persistence&split=test&horizon=60]/detected_onsets | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.pev.false_alarm_episodes | Persistence 60-min warning (fires at 150 mg/dL), final test: false_alarm_episodes | 70 | count | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /events[model=persistence&split=test&horizon=60]/false_alarm_episodes | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.pev.false_alarms_per_observed_patient_day | Persistence 60-min warning (fires at 150 mg/dL), final test: false_alarms_per_observed_patient_day | 2.6114 | count | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /events[model=persistence&split=test&horizon=60]/false_alarms_per_observed_patient_day | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.iv.coverage30 | Served 80 % range, final test, 30 min: coverage | 0.74204 | share | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /intervals[label=lightgbm:dyn-prof-v2&personalization=adaptive&split=test&horizon=30]/coverage | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.iv.width30 | Served 80 % range, final test, 30 min: width | 33.4234 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /intervals[label=lightgbm:dyn-prof-v2&personalization=adaptive&split=test&horizon=30]/width | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.iv.q30 | Served 80 % range, final test, 30 min: q | 16.7117 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /intervals[label=lightgbm:dyn-prof-v2&personalization=adaptive&split=test&horizon=30]/q | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.iv.per_patient_min30 | Served 80 % range, final test, 30 min: per_patient_min | 0.60264 | share | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /intervals[label=lightgbm:dyn-prof-v2&personalization=adaptive&split=test&horizon=30]/per_patient_min | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.iv.per_patient_max30 | Served 80 % range, final test, 30 min: per_patient_max | 0.8231 | share | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /intervals[label=lightgbm:dyn-prof-v2&personalization=adaptive&split=test&horizon=30]/per_patient_max | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.iv.n30 | Served 80 % range, final test, 30 min: n | 7699 | origins | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /intervals[label=lightgbm:dyn-prof-v2&personalization=adaptive&split=test&horizon=30]/n | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.iv.coverage60 | Served 80 % range, final test, 60 min: coverage | 0.73979 | share | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /intervals[label=lightgbm:dyn-prof-v2&personalization=adaptive&split=test&horizon=60]/coverage | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.iv.width60 | Served 80 % range, final test, 60 min: width | 52.7294 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /intervals[label=lightgbm:dyn-prof-v2&personalization=adaptive&split=test&horizon=60]/width | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.iv.q60 | Served 80 % range, final test, 60 min: q | 26.3647 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /intervals[label=lightgbm:dyn-prof-v2&personalization=adaptive&split=test&horizon=60]/q | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.iv.per_patient_min60 | Served 80 % range, final test, 60 min: per_patient_min | 0.51429 | share | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /intervals[label=lightgbm:dyn-prof-v2&personalization=adaptive&split=test&horizon=60]/per_patient_min | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.iv.per_patient_max60 | Served 80 % range, final test, 60 min: per_patient_max | 0.87568 | share | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /intervals[label=lightgbm:dyn-prof-v2&personalization=adaptive&split=test&horizon=60]/per_patient_max | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.iv.n60 | Served 80 % range, final test, 60 min: n | 7690 | origins | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /intervals[label=lightgbm:dyn-prof-v2&personalization=adaptive&split=test&horizon=60]/n | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.training.origins | Cohort, training segment: origins | 23745 | origins | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/training/origins | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.training.issued_forecast_origins | Cohort, training segment: issued_forecast_origins | 23577 | origins | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/training/issued_forecast_origins | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.training.covered_hours | Cohort, training segment: covered_hours | 1978.8 | hours | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/training/covered_hours | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.training.matched_targets_30 | Cohort, training segment: matched_targets_30 | 23567 | targets | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/training/matched_targets_30 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.training.matched_targets_60 | Cohort, training segment: matched_targets_60 | 23567 | targets | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/training/matched_targets_60 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.training.distinct_new_onsets_60 | Cohort, training segment: distinct_new_onsets_60 | 203 | onsets | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/training/distinct_new_onsets_60 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.validation.origins | Cohort, validation segment: origins | 3766 | origins | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/validation/origins | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.validation.issued_forecast_origins | Cohort, validation segment: issued_forecast_origins | 3766 | origins | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/validation/issued_forecast_origins | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.validation.covered_hours | Cohort, validation segment: covered_hours | 313.8 | hours | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/validation/covered_hours | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.validation.matched_targets_30 | Cohort, validation segment: matched_targets_30 | 3766 | targets | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/validation/matched_targets_30 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.validation.matched_targets_60 | Cohort, validation segment: matched_targets_60 | 3766 | targets | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/validation/matched_targets_60 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.validation.distinct_new_onsets_60 | Cohort, validation segment: distinct_new_onsets_60 | 32 | onsets | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/validation/distinct_new_onsets_60 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.calibration.origins | Cohort, calibration segment: origins | 3766 | origins | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/calibration/origins | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.calibration.issued_forecast_origins | Cohort, calibration segment: issued_forecast_origins | 3766 | origins | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/calibration/issued_forecast_origins | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.calibration.covered_hours | Cohort, calibration segment: covered_hours | 313.8 | hours | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/calibration/covered_hours | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.calibration.matched_targets_30 | Cohort, calibration segment: matched_targets_30 | 3766 | targets | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/calibration/matched_targets_30 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.calibration.matched_targets_60 | Cohort, calibration segment: matched_targets_60 | 3766 | targets | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/calibration/matched_targets_60 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.calibration.distinct_new_onsets_60 | Cohort, calibration segment: distinct_new_onsets_60 | 28 | onsets | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/calibration/distinct_new_onsets_60 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.test.origins | Cohort, test segment: origins | 7730 | origins | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/test/origins | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.test.issued_forecast_origins | Cohort, test segment: issued_forecast_origins | 7720 | origins | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/test/issued_forecast_origins | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.test.covered_hours | Cohort, test segment: covered_hours | 644.2 | hours | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/test/covered_hours | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.test.matched_targets_30 | Cohort, test segment: matched_targets_30 | 7699 | targets | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/test/matched_targets_30 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.test.matched_targets_60 | Cohort, test segment: matched_targets_60 | 7690 | targets | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/test/matched_targets_60 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.seg.test.distinct_new_onsets_60 | Cohort, test segment: distinct_new_onsets_60 | 64 | onsets | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/by_segment/test/distinct_new_onsets_60 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.participants | Participants in the serving experiment (Type 2 group) | 14 | people | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /cohort/participants | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.excluded | CGMacros participants outside the Type 2 group (length of /excluded) | 31 | people | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /excluded | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.sens.meal60 | Optimistic meal-logging sensitivity (not eligible for serving), test, 60 min | 19.0731 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /sensitivity[experiment=optimistic meal logging&split=test&horizon=60]/mae | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.model_version | Model version evaluated | lightgbm.dyn-prof-v2.gt-primary-2026-10-13-fusion | text | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /model_version | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.created_at | Experiment run time | 2026-10-13 08:02 | UTC | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /created_at | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.selection_rule | Declared selection rule | Lowest equal-weight mean of per-patient MAE across the 30- and 60-minute horizons on the validation-selection segment, population model without patient adaptation. | text | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/rule | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.better30 | Serving test: participants with lower serving error than persistence, 30 min (derived count) | 14 | people | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /per_patient_test (count of people with serving MAE < persistence MAE, 30 min) | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
fus.better60 | Serving test: participants with lower serving error than persistence, 60 min (derived count) | 13 | people | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /per_patient_test (count of people with serving MAE < persistence MAE, 60 min) | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
wu.onsets | Serving 60-min warning, final test: eligible onsets (denominator of every notice share) | 64 | onsets | fusion_pred | artifacts/experiments/gt-primary-2026-10-13-fusion/predictions.parquet | scripts/warning_utility.py demo row (WT.demo, key=demo) → onsets | 2026-10-13T08:02:25+00:00 (file time) | 1a5a7164fd8d17f36a4248a5a9f30f2dfee53b30b010ca62fab30df8c77b1778 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
wu.detected | Serving 60-min warning, final test: detected onsets | 61 | onsets | fusion_pred | artifacts/experiments/gt-primary-2026-10-13-fusion/predictions.parquet | scripts/warning_utility.py demo row (WT.demo, key=demo) → detected | 2026-10-13T08:02:25+00:00 (file time) | 1a5a7164fd8d17f36a4248a5a9f30f2dfee53b30b010ca62fab30df8c77b1778 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
wu.ge15 | Serving 60-min warning, final test: onsets warned ≥ 15 min ahead (qualified lead) | 43 | onsets | fusion_pred | artifacts/experiments/gt-primary-2026-10-13-fusion/predictions.parquet | scripts/warning_utility.py demo row (WT.demo, key=demo) → ge15 | 2026-10-13T08:02:25+00:00 (file time) | 1a5a7164fd8d17f36a4248a5a9f30f2dfee53b30b010ca62fab30df8c77b1778 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
wu.ge30 | Serving 60-min warning, final test: onsets warned ≥ 30 min ahead (qualified lead) | 31 | onsets | fusion_pred | artifacts/experiments/gt-primary-2026-10-13-fusion/predictions.parquet | scripts/warning_utility.py demo row (WT.demo, key=demo) → ge30 | 2026-10-13T08:02:25+00:00 (file time) | 1a5a7164fd8d17f36a4248a5a9f30f2dfee53b30b010ca62fab30df8c77b1778 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
wu.ge60 | Serving 60-min warning, final test: onsets warned ≥ 60 min ahead (qualified lead) | 13 | onsets | fusion_pred | artifacts/experiments/gt-primary-2026-10-13-fusion/predictions.parquet | scripts/warning_utility.py demo row (WT.demo, key=demo) → ge60 | 2026-10-13T08:02:25+00:00 (file time) | 1a5a7164fd8d17f36a4248a5a9f30f2dfee53b30b010ca62fab30df8c77b1778 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
wu.firing | Serving 60-min warning, final test: eligible origins at which the warning was active | 596 | origins | fusion_pred | artifacts/experiments/gt-primary-2026-10-13-fusion/predictions.parquet | scripts/warning_utility.py demo row (WT.demo, key=demo) → firing | 2026-10-13T08:02:25+00:00 (file time) | 1a5a7164fd8d17f36a4248a5a9f30f2dfee53b30b010ca62fab30df8c77b1778 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
wu.eligible | Serving 60-min warning, final test: eligible origins (time-in-warning denominator) | 5042 | origins | fusion_pred | artifacts/experiments/gt-primary-2026-10-13-fusion/predictions.parquet | scripts/warning_utility.py demo row (WT.demo, key=demo) → eligible | 2026-10-13T08:02:25+00:00 (file time) | 1a5a7164fd8d17f36a4248a5a9f30f2dfee53b30b010ca62fab30df8c77b1778 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
wu.share_eligible | Serving 60-min warning, final test: share of eligible origins in warning | 0.11821 | share | fusion_pred | artifacts/experiments/gt-primary-2026-10-13-fusion/predictions.parquet | scripts/warning_utility.py demo row (WT.demo, key=demo) → share_eligible | 2026-10-13T08:02:25+00:00 (file time) | 1a5a7164fd8d17f36a4248a5a9f30f2dfee53b30b010ca62fab30df8c77b1778 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
wu.lead_median | Serving 60-min warning, final test: median qualified lead (detected onsets) | 30 | min | fusion_pred | artifacts/experiments/gt-primary-2026-10-13-fusion/predictions.parquet | scripts/warning_utility.py demo row (WT.demo, key=demo) → lead_median | 2026-10-13T08:02:25+00:00 (file time) | 1a5a7164fd8d17f36a4248a5a9f30f2dfee53b30b010ca62fab30df8c77b1778 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
wu.n_leads | Serving 60-min warning, final test: number of qualified leads (detected onsets) | 61 | onsets | fusion_pred | artifacts/experiments/gt-primary-2026-10-13-fusion/predictions.parquet | scripts/warning_utility.py demo row (WT.demo, key=demo) → n_leads | 2026-10-13T08:02:25+00:00 (file time) | 1a5a7164fd8d17f36a4248a5a9f30f2dfee53b30b010ca62fab30df8c77b1778 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
wu.lead_q25 | Serving 60-min warning, final test: qualified lead 25th percentile | 10 | min | fusion_pred | artifacts/experiments/gt-primary-2026-10-13-fusion/predictions.parquet | scripts/warning_utility.py demo row (WT.demo, key=demo) → lead_q25 | 2026-10-13T08:02:25+00:00 (file time) | 1a5a7164fd8d17f36a4248a5a9f30f2dfee53b30b010ca62fab30df8c77b1778 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
wu.lead_q75 | Serving 60-min warning, final test: qualified lead 75th percentile | 55 | min | fusion_pred | artifacts/experiments/gt-primary-2026-10-13-fusion/predictions.parquet | scripts/warning_utility.py demo row (WT.demo, key=demo) → lead_q75 | 2026-10-13T08:02:25+00:00 (file time) | 1a5a7164fd8d17f36a4248a5a9f30f2dfee53b30b010ca62fab30df8c77b1778 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
tr.all.mae30 | Validation segment, all, 30 min: macro MAE | 13.3585 | mg/dL | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=30&trend=all]/mae_macro | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.all.cov30 | Validation segment, all, 30 min: served range coverage | 0.7273 | share | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=30&trend=all]/range_coverage | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.all.n30 | Validation segment, all, 30 min: origins | 3766 | origins | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=30&trend=all]/origins | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.falling_fast.mae30 | Validation segment, falling_fast, 30 min: macro MAE | 18.7831 | mg/dL | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=30&trend=falling_fast]/mae_macro | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.falling_fast.cov30 | Validation segment, falling_fast, 30 min: served range coverage | 0.52 | share | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=30&trend=falling_fast]/range_coverage | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.falling_fast.n30 | Validation segment, falling_fast, 30 min: origins | 100 | origins | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=30&trend=falling_fast]/origins | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.falling.mae30 | Validation segment, falling, 30 min: macro MAE | 15.2895 | mg/dL | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=30&trend=falling]/mae_macro | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.falling.cov30 | Validation segment, falling, 30 min: served range coverage | 0.64458 | share | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=30&trend=falling]/range_coverage | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.falling.n30 | Validation segment, falling, 30 min: origins | 332 | origins | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=30&trend=falling]/origins | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.steady.mae30 | Validation segment, steady, 30 min: macro MAE | 11.9267 | mg/dL | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=30&trend=steady]/mae_macro | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.steady.cov30 | Validation segment, steady, 30 min: served range coverage | 0.77299 | share | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=30&trend=steady]/range_coverage | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.steady.n30 | Validation segment, steady, 30 min: origins | 2947 | origins | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=30&trend=steady]/origins | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.rising.mae30 | Validation segment, rising, 30 min: macro MAE | 18.9525 | mg/dL | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=30&trend=rising]/mae_macro | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.rising.cov30 | Validation segment, rising, 30 min: served range coverage | 0.56225 | share | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=30&trend=rising]/range_coverage | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.rising.n30 | Validation segment, rising, 30 min: origins | 249 | origins | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=30&trend=rising]/origins | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.rising_fast.mae30 | Validation segment, rising_fast, 30 min: macro MAE | 25.8734 | mg/dL | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=30&trend=rising_fast]/mae_macro | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.rising_fast.cov30 | Validation segment, rising_fast, 30 min: served range coverage | 0.39855 | share | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=30&trend=rising_fast]/range_coverage | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.rising_fast.n30 | Validation segment, rising_fast, 30 min: origins | 138 | origins | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=30&trend=rising_fast]/origins | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.bias_high30 | Validation segment, 30 min: mean signed error (forecast − actual) when actual > 180 | -7.7071 | mg/dL | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=30&trend=all]/mean_signed_error_target_above_180 | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.all.mae60 | Validation segment, all, 60 min: macro MAE | 20.9193 | mg/dL | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=60&trend=all]/mae_macro | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.all.cov60 | Validation segment, all, 60 min: served range coverage | 0.72756 | share | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=60&trend=all]/range_coverage | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.all.n60 | Validation segment, all, 60 min: origins | 3766 | origins | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=60&trend=all]/origins | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.falling_fast.mae60 | Validation segment, falling_fast, 60 min: macro MAE | 23.4836 | mg/dL | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=60&trend=falling_fast]/mae_macro | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.falling_fast.cov60 | Validation segment, falling_fast, 60 min: served range coverage | 0.61 | share | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=60&trend=falling_fast]/range_coverage | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.falling_fast.n60 | Validation segment, falling_fast, 60 min: origins | 100 | origins | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=60&trend=falling_fast]/origins | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.falling.mae60 | Validation segment, falling, 60 min: macro MAE | 21.811 | mg/dL | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=60&trend=falling]/mae_macro | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.falling.cov60 | Validation segment, falling, 60 min: served range coverage | 0.64458 | share | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=60&trend=falling]/range_coverage | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.falling.n60 | Validation segment, falling, 60 min: origins | 332 | origins | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=60&trend=falling]/origins | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.steady.mae60 | Validation segment, steady, 60 min: macro MAE | 19.536 | mg/dL | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=60&trend=steady]/mae_macro | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.steady.cov60 | Validation segment, steady, 60 min: served range coverage | 0.77027 | share | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=60&trend=steady]/range_coverage | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.steady.n60 | Validation segment, steady, 60 min: origins | 2947 | origins | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=60&trend=steady]/origins | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.rising.mae60 | Validation segment, rising, 60 min: macro MAE | 28.7608 | mg/dL | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=60&trend=rising]/mae_macro | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.rising.cov60 | Validation segment, rising, 60 min: served range coverage | 0.53012 | share | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=60&trend=rising]/range_coverage | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.rising.n60 | Validation segment, rising, 60 min: origins | 249 | origins | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=60&trend=rising]/origins | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.rising_fast.mae60 | Validation segment, rising_fast, 60 min: macro MAE | 35.3262 | mg/dL | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=60&trend=rising_fast]/mae_macro | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.rising_fast.cov60 | Validation segment, rising_fast, 60 min: served range coverage | 0.45652 | share | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=60&trend=rising_fast]/range_coverage | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.rising_fast.n60 | Validation segment, rising_fast, 60 min: origins | 138 | origins | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=60&trend=rising_fast]/origins | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
tr.bias_high60 | Validation segment, 60 min: mean signed error (forecast − actual) when actual > 180 | -16.9194 | mg/dL | trend | artifacts/experiments/gt-primary-2026-10-13-fusion/error_by_trend.json | /rows[horizon=60&trend=all]/mean_signed_error_target_above_180 | 2026-10-20T17:41:13+00:00 | 5c9d1104dee3a8b8b5e94771c2e92b94f16465cb19ceacc791d6db065a976c24 | scripts/error_by_trend.py | local / reproducible-only (git-ignored)
parity.compared | Saved app-vs-evaluation parity run: compared | 15440 | forecasts | parity | reports/serving_parity.json | /compared | 2026-10-20T17:44:10+00:00 | d43b660ed9cc0189f9f5b34b6e997ce700339ddea80c1a60657519b06ea49664 | scripts/serving_parity.py | public in repo (tracked)
parity.offline_forecasts | Saved app-vs-evaluation parity run: offline_forecasts | 15440 | forecasts | parity | reports/serving_parity.json | /offline_forecasts | 2026-10-20T17:44:10+00:00 | d43b660ed9cc0189f9f5b34b6e997ce700339ddea80c1a60657519b06ea49664 | scripts/serving_parity.py | public in repo (tracked)
parity.missing_online | Saved app-vs-evaluation parity run: missing_online | 0 | forecasts | parity | reports/serving_parity.json | /missing_online | 2026-10-20T17:44:10+00:00 | d43b660ed9cc0189f9f5b34b6e997ce700339ddea80c1a60657519b06ea49664 | scripts/serving_parity.py | public in repo (tracked)
parity.over_tolerance | Saved app-vs-evaluation parity run: over_tolerance | 0 | forecasts | parity | reports/serving_parity.json | /over_tolerance | 2026-10-20T17:44:10+00:00 | d43b660ed9cc0189f9f5b34b6e997ce700339ddea80c1a60657519b06ea49664 | scripts/serving_parity.py | public in repo (tracked)
parity.tolerance_mg_dl | Saved app-vs-evaluation parity run: tolerance_mg_dl | 0.05 | mg/dL | parity | reports/serving_parity.json | /tolerance_mg_dl | 2026-10-20T17:44:10+00:00 | d43b660ed9cc0189f9f5b34b6e997ce700339ddea80c1a60657519b06ea49664 | scripts/serving_parity.py | public in repo (tracked)
parity.max_abs_diff_mg_dl | Saved app-vs-evaluation parity run: max_abs_diff_mg_dl | 0.049998 | mg/dL | parity | reports/serving_parity.json | /max_abs_diff_mg_dl | 2026-10-20T17:44:10+00:00 | d43b660ed9cc0189f9f5b34b6e997ce700339ddea80c1a60657519b06ea49664 | scripts/serving_parity.py | public in repo (tracked)
parity.participants | Saved app-vs-evaluation parity run: participants | 14 | people | parity | reports/serving_parity.json | /participants | 2026-10-20T17:44:10+00:00 | d43b660ed9cc0189f9f5b34b6e997ce700339ddea80c1a60657519b06ea49664 | scripts/serving_parity.py | public in repo (tracked)
parity.outcome | Saved app-vs-evaluation parity run: outcome | pass | text | parity | reports/serving_parity.json | /outcome | 2026-10-20T17:44:10+00:00 | d43b660ed9cc0189f9f5b34b6e997ce700339ddea80c1a60657519b06ea49664 | scripts/serving_parity.py | public in repo (tracked)
parity.model_version | Saved app-vs-evaluation parity run: model_version | lightgbm.dyn-prof-v2.gt-primary-2026-10-13-fusion | text | parity | reports/serving_parity.json | /model_version | 2026-10-20T17:44:10+00:00 | d43b660ed9cc0189f9f5b34b6e997ce700339ddea80c1a60657519b06ea49664 | scripts/serving_parity.py | public in repo (tracked)
parity_final.compared | Parity re-run in the final verification ledger: compared | 15440 | forecasts | parity_final | reports/verification/final_2026-10-22_145326Z/parity/serving_parity.json | /compared | 2026-10-22T15:02:11+00:00 | 390c76b31f1d8546289b6dcbdca99e37c0ebe721e4b28199783060e50e1de489 | scripts/serving_parity.py | public in repo (tracked)
parity_final.offline_forecasts | Parity re-run in the final verification ledger: offline_forecasts | 15440 | forecasts | parity_final | reports/verification/final_2026-10-22_145326Z/parity/serving_parity.json | /offline_forecasts | 2026-10-22T15:02:11+00:00 | 390c76b31f1d8546289b6dcbdca99e37c0ebe721e4b28199783060e50e1de489 | scripts/serving_parity.py | public in repo (tracked)
parity_final.missing_online | Parity re-run in the final verification ledger: missing_online | 0 | forecasts | parity_final | reports/verification/final_2026-10-22_145326Z/parity/serving_parity.json | /missing_online | 2026-10-22T15:02:11+00:00 | 390c76b31f1d8546289b6dcbdca99e37c0ebe721e4b28199783060e50e1de489 | scripts/serving_parity.py | public in repo (tracked)
parity_final.over_tolerance | Parity re-run in the final verification ledger: over_tolerance | 0 | forecasts | parity_final | reports/verification/final_2026-10-22_145326Z/parity/serving_parity.json | /over_tolerance | 2026-10-22T15:02:11+00:00 | 390c76b31f1d8546289b6dcbdca99e37c0ebe721e4b28199783060e50e1de489 | scripts/serving_parity.py | public in repo (tracked)
parity_final.tolerance_mg_dl | Parity re-run in the final verification ledger: tolerance_mg_dl | 0.05 | mg/dL | parity_final | reports/verification/final_2026-10-22_145326Z/parity/serving_parity.json | /tolerance_mg_dl | 2026-10-22T15:02:11+00:00 | 390c76b31f1d8546289b6dcbdca99e37c0ebe721e4b28199783060e50e1de489 | scripts/serving_parity.py | public in repo (tracked)
parity_final.max_abs_diff_mg_dl | Parity re-run in the final verification ledger: max_abs_diff_mg_dl | 0.049998 | mg/dL | parity_final | reports/verification/final_2026-10-22_145326Z/parity/serving_parity.json | /max_abs_diff_mg_dl | 2026-10-22T15:02:11+00:00 | 390c76b31f1d8546289b6dcbdca99e37c0ebe721e4b28199783060e50e1de489 | scripts/serving_parity.py | public in repo (tracked)
parity_final.participants | Parity re-run in the final verification ledger: participants | 14 | people | parity_final | reports/verification/final_2026-10-22_145326Z/parity/serving_parity.json | /participants | 2026-10-22T15:02:11+00:00 | 390c76b31f1d8546289b6dcbdca99e37c0ebe721e4b28199783060e50e1de489 | scripts/serving_parity.py | public in repo (tracked)
parity_final.outcome | Parity re-run in the final verification ledger: outcome | pass | text | parity_final | reports/verification/final_2026-10-22_145326Z/parity/serving_parity.json | /outcome | 2026-10-22T15:02:11+00:00 | 390c76b31f1d8546289b6dcbdca99e37c0ebe721e4b28199783060e50e1de489 | scripts/serving_parity.py | public in repo (tracked)
parity_final.model_version | Parity re-run in the final verification ledger: model_version | lightgbm.dyn-prof-v2.gt-primary-2026-10-13-fusion | text | parity_final | reports/verification/final_2026-10-22_145326Z/parity/serving_parity.json | /model_version | 2026-10-22T15:02:11+00:00 | 390c76b31f1d8546289b6dcbdca99e37c0ebe721e4b28199783060e50e1de489 | scripts/serving_parity.py | public in repo (tracked)
nest.sel.mae30 | Nested internal (14 people held out once): selected final, 30 min MAE | 13.884 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /tracks/multimodal_5min/pipelines[pipeline=selected&mode=final]/horizons/30/mae | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
nest.sel.mae60 | Nested internal (14 people held out once): selected final, 60 min MAE | 22.305 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /tracks/multimodal_5min/pipelines[pipeline=selected&mode=final]/horizons/60/mae | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
nest.pers.mae30 | Nested internal (14 people held out once): persistence:- off, 30 min MAE | 15.339 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /tracks/multimodal_5min/pipelines[pipeline=persistence:-&mode=off]/horizons/30/mae | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
nest.pers.mae60 | Nested internal (14 people held out once): persistence:- off, 60 min MAE | 24.467 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /tracks/multimodal_5min/pipelines[pipeline=persistence:-&mode=off]/horizons/60/mae | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
nest.v2.mae30 | Nested internal (14 people held out once): v2_recipe:dyn-v2 v2_policy, 30 min MAE | 14.52 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /tracks/multimodal_5min/pipelines[pipeline=v2_recipe:dyn-v2&mode=v2_policy]/horizons/30/mae | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
nest.v2.mae60 | Nested internal (14 people held out once): v2_recipe:dyn-v2 v2_policy, 60 min MAE | 24.253 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /tracks/multimodal_5min/pipelines[pipeline=v2_recipe:dyn-v2&mode=v2_policy]/horizons/60/mae | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
nest.cov30 | Nested internal: selected procedure 80 % range coverage, 30 min | 0.80074 | share | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /tracks/multimodal_5min/pipelines[pipeline=selected&mode=final]/intervals/30/global/coverage_pooled | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
nest.prof.d30 | Nested: profile minus no profile (LightGBM L2, population), 30 min | 0.335 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /tracks/multimodal_5min/comparisons[comparison=Profile inputs: LightGBM L2 dyn-prof-v2 minus dyn-v2&horizon=30]/mean_diff | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
nest.prof.lo30 | 95 % CI lower | -0.275 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /tracks/multimodal_5min/comparisons[comparison=Profile inputs: LightGBM L2 dyn-prof-v2 minus dyn-v2&horizon=30]/ci95[0] | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
nest.prof.hi30 | 95 % CI upper | 0.894 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /tracks/multimodal_5min/comparisons[comparison=Profile inputs: LightGBM L2 dyn-prof-v2 minus dyn-v2&horizon=30]/ci95[1] | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
ext.cand.mae30 | External ShanghaiT2DM (98): 15-min LightGBM, 30 min MAE | 11.153 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /external/pipelines[pipeline=lightgbm_l2:cgm15-v1&mode=final]/horizons/30/mae | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
ext.pers.mae30 | External ShanghaiT2DM: persistence, 30 min MAE | 12.145 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /external/pipelines[pipeline=persistence:-&mode=off]/horizons/30/mae | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
ext.cmp.mean_diff30 | External: candidate minus persistence, 30 min, mean_diff | -0.992 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /external/comparisons[comparison=Selected candidate minus persistence&horizon=30]/mean_diff | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
ext.cmp.lo30 | External: candidate minus persistence, 30 min, ci95[0] | -1.211 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /external/comparisons[comparison=Selected candidate minus persistence&horizon=30]/ci95[0] | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
ext.cmp.hi30 | External: candidate minus persistence, 30 min, ci95[1] | -0.766 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /external/comparisons[comparison=Selected candidate minus persistence&horizon=30]/ci95[1] | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
ext.cmp.improved30 | External: candidate minus persistence, 30 min, improved | 80 | people | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /external/comparisons[comparison=Selected candidate minus persistence&horizon=30]/improved | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
ext.cmp.n_participants30 | External: candidate minus persistence, 30 min, n_participants | 98 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /external/comparisons[comparison=Selected candidate minus persistence&horizon=30]/n_participants | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
ext.cov30 | External: 80 % range coverage, 30 min | 0.87057 | share | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /external/pipelines[pipeline=lightgbm_l2:cgm15-v1&mode=final]/intervals/30/global/coverage_pooled | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
nest.cov60 | Nested internal: selected procedure 80 % range coverage, 60 min | 0.78811 | share | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /tracks/multimodal_5min/pipelines[pipeline=selected&mode=final]/intervals/60/global/coverage_pooled | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
nest.prof.d60 | Nested: profile minus no profile (LightGBM L2, population), 60 min | -0.747 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /tracks/multimodal_5min/comparisons[comparison=Profile inputs: LightGBM L2 dyn-prof-v2 minus dyn-v2&horizon=60]/mean_diff | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
nest.prof.lo60 | 95 % CI lower | -2.187 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /tracks/multimodal_5min/comparisons[comparison=Profile inputs: LightGBM L2 dyn-prof-v2 minus dyn-v2&horizon=60]/ci95[0] | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
nest.prof.hi60 | 95 % CI upper | 0.608 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /tracks/multimodal_5min/comparisons[comparison=Profile inputs: LightGBM L2 dyn-prof-v2 minus dyn-v2&horizon=60]/ci95[1] | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
ext.cand.mae60 | External ShanghaiT2DM (98): 15-min LightGBM, 60 min MAE | 20.105 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /external/pipelines[pipeline=lightgbm_l2:cgm15-v1&mode=final]/horizons/60/mae | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
ext.pers.mae60 | External ShanghaiT2DM: persistence, 60 min MAE | 20.444 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /external/pipelines[pipeline=persistence:-&mode=off]/horizons/60/mae | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
ext.cmp.mean_diff60 | External: candidate minus persistence, 60 min, mean_diff | -0.339 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /external/comparisons[comparison=Selected candidate minus persistence&horizon=60]/mean_diff | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
ext.cmp.lo60 | External: candidate minus persistence, 60 min, ci95[0] | -0.762 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /external/comparisons[comparison=Selected candidate minus persistence&horizon=60]/ci95[0] | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
ext.cmp.hi60 | External: candidate minus persistence, 60 min, ci95[1] | 0.1 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /external/comparisons[comparison=Selected candidate minus persistence&horizon=60]/ci95[1] | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
ext.cmp.improved60 | External: candidate minus persistence, 60 min, improved | 66 | people | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /external/comparisons[comparison=Selected candidate minus persistence&horizon=60]/improved | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
ext.cmp.n_participants60 | External: candidate minus persistence, 60 min, n_participants | 98 | mg/dL | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /external/comparisons[comparison=Selected candidate minus persistence&horizon=60]/n_participants | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
ext.cov60 | External: 80 % range coverage, 60 min | 0.86063 | share | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /external/pipelines[pipeline=lightgbm_l2:cgm15-v1&mode=final]/intervals/60/global/coverage_pooled | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
nest.ridge_folds | Nested: outer folds whose rule chose Ridge CGM-only | 4 | folds | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /tracks/multimodal_5min/selection_by_fold (count of outer folds selecting ridge:cgm-only-v1) | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
nest.frozen_at | Nested protocol frozen | 2026-10-12 13:29 | UTC | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /frozen_at | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
nest.opened_at | Nested internal assessment opened | 2026-10-12 13:43 | UTC | nested | artifacts/research/gt-nested-2026-10-12-v1/report.json | /assessment_opened/opened_at | 2026-10-12T13:28:51+00:00 | 368c28117b15d8092b555cf31176aabdd9d8f3bc3bf5137097ef1ce6fec6466b | protocol configs/experiments/gt-nested-2026-10-12-v1.yaml (first commit 90121a4 2026-10-12T20:39:33+05:30) | local / reproducible-only (git-ignored)
2h.sel.mae30 | Two-hour study: selected personalised, 30 min MAE | 13.953 | mg/dL | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /forecasts/selected¦final/horizons/30/mae | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.pers.mae30 | Two-hour study: persistence, 30 min MAE | 15.339 | mg/dL | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /forecasts/persistence:-¦off/horizons/30/mae | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.sel.mae60 | Two-hour study: selected personalised, 60 min MAE | 22.288 | mg/dL | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /forecasts/selected¦final/horizons/60/mae | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.pers.mae60 | Two-hour study: persistence, 60 min MAE | 24.467 | mg/dL | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /forecasts/persistence:-¦off/horizons/60/mae | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.sel.mae120 | Two-hour study: selected personalised, 120 min MAE | 30.617 | mg/dL | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /forecasts/selected¦final/horizons/120/mae | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.pers.mae120 | Two-hour study: persistence, 120 min MAE | 35.869 | mg/dL | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /forecasts/persistence:-¦off/horizons/120/mae | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.cmp.d | Two-hour: selected minus persistence at 120 min | -5.25 | mg/dL | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /comparisons[comparison=Selected (personalised) minus persistence&horizon=120]/mean_difference | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.cmp.lo | 95 % CI lower | -6.44 | mg/dL | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /comparisons[comparison=Selected (personalised) minus persistence&horizon=120]/ci95[0] | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.cmp.hi | 95 % CI upper | -4.05 | mg/dL | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /comparisons[comparison=Selected (personalised) minus persistence&horizon=120]/ci95[1] | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.cmp.better | People with lower error than persistence | 14 | people | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /comparisons[comparison=Selected (personalised) minus persistence&horizon=120]/better_for | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.cmp.n | Participants | 14 | people | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /comparisons[comparison=Selected (personalised) minus persistence&horizon=120]/participants | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.prof.d | Two-hour profile ablation (LightGBM), 120 min | -0.86 | mg/dL | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /comparisons[comparison=Profile (EHR-like) inputs: LightGBM dyn-prof-v2 minus dyn-v2&horizon=120]/mean_difference | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.prof.lo | 95 % CI lower | -2.6 | mg/dL | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /comparisons[comparison=Profile (EHR-like) inputs: LightGBM dyn-prof-v2 minus dyn-v2&horizon=120]/ci95[0] | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.prof.hi | 95 % CI upper | 1.02 | mg/dL | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /comparisons[comparison=Profile (EHR-like) inputs: LightGBM dyn-prof-v2 minus dyn-v2&horizon=120]/ci95[1] | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.w.true_onsets | Two-hour 120-min warning: true_onsets | 329 | onsets | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /warnings[pipeline=selected&mode=final&horizon=120]/true_onsets | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.w.detected_onsets | Two-hour 120-min warning: detected_onsets | 315 | onsets | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /warnings[pipeline=selected&mode=final&horizon=120]/detected_onsets | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.w.false_alarms_per_observed_day | Two-hour 120-min warning: false_alarms_per_observed_day | 1.623 | per observed day | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /warnings[pipeline=selected&mode=final&horizon=120]/false_alarms_per_observed_day | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.w.qualified_lead_median | Two-hour 120-min warning: qualified_lead_median | 15 | min | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /warnings[pipeline=selected&mode=final&horizon=120]/qualified_lead_median | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.w.qualified_lead_iqr_lo | Two-hour 120-min warning: qualified_lead_iqr[0] | 5 | min | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /warnings[pipeline=selected&mode=final&horizon=120]/qualified_lead_iqr[0] | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.w.qualified_lead_iqr_hi | Two-hour 120-min warning: qualified_lead_iqr[1] | 65 | min | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /warnings[pipeline=selected&mode=final&horizon=120]/qualified_lead_iqr[1] | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.cov120 | Two-hour: 80 % range coverage at 120 min | 0.76488 | share | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /forecasts/selected¦final/coverage/120 | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.frozen_at | Two-hour protocol frozen | 2026-10-20 09:18 | UTC | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /frozen_at | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2h.opened_at | Two-hour assessment opened | 2026-10-20 09:36 | UTC | twoh | artifacts/research/gt-2h-2026-10-20/two_hour_summary.json | /assessment_opened/opened_at | 2026-10-20T09:18:07+00:00 | 5d09f6be55284246f450ae2093a700162aa9c0a6cbfcf3fbd3d7f65191fdc8ed | protocol configs/experiments/gt-2h-2026-10-20.yaml (first commit 4d66df4 2026-10-20T14:50:24+05:30) | local / reproducible-only (git-ignored)
2hx.mae30 | External two-hour (98 ShanghaiT2DM): 30 min MAE | 10.898 | mg/dL | twoh_ext | artifacts/research/gt-2h-ext-2026-10-20/two_hour_external_summary.json | /mae/30 | 2026-10-20T10:10:25+00:00 | 58a1ba2570f9fe0026d918d914a11fc508b5083d03ce357bcecfe500220633ee | protocol configs/experiments/gt-2h-ext-2026-10-20.yaml (first commit c55789a 2026-10-20T15:38:49+05:30) | local / reproducible-only (git-ignored)
2hx.pers30 | External two-hour: persistence 30 min MAE | 12.145 | mg/dL | twoh_ext | artifacts/research/gt-2h-ext-2026-10-20/two_hour_external_summary.json | /persistence_mae/30 | 2026-10-20T10:10:25+00:00 | 58a1ba2570f9fe0026d918d914a11fc508b5083d03ce357bcecfe500220633ee | protocol configs/experiments/gt-2h-ext-2026-10-20.yaml (first commit c55789a 2026-10-20T15:38:49+05:30) | local / reproducible-only (git-ignored)
2hx.mae60 | External two-hour (98 ShanghaiT2DM): 60 min MAE | 19.311 | mg/dL | twoh_ext | artifacts/research/gt-2h-ext-2026-10-20/two_hour_external_summary.json | /mae/60 | 2026-10-20T10:10:25+00:00 | 58a1ba2570f9fe0026d918d914a11fc508b5083d03ce357bcecfe500220633ee | protocol configs/experiments/gt-2h-ext-2026-10-20.yaml (first commit c55789a 2026-10-20T15:38:49+05:30) | local / reproducible-only (git-ignored)
2hx.pers60 | External two-hour: persistence 60 min MAE | 20.444 | mg/dL | twoh_ext | artifacts/research/gt-2h-ext-2026-10-20/two_hour_external_summary.json | /persistence_mae/60 | 2026-10-20T10:10:25+00:00 | 58a1ba2570f9fe0026d918d914a11fc508b5083d03ce357bcecfe500220633ee | protocol configs/experiments/gt-2h-ext-2026-10-20.yaml (first commit c55789a 2026-10-20T15:38:49+05:30) | local / reproducible-only (git-ignored)
2hx.mae120 | External two-hour (98 ShanghaiT2DM): 120 min MAE | 29.553 | mg/dL | twoh_ext | artifacts/research/gt-2h-ext-2026-10-20/two_hour_external_summary.json | /mae/120 | 2026-10-20T10:10:25+00:00 | 58a1ba2570f9fe0026d918d914a11fc508b5083d03ce357bcecfe500220633ee | protocol configs/experiments/gt-2h-ext-2026-10-20.yaml (first commit c55789a 2026-10-20T15:38:49+05:30) | local / reproducible-only (git-ignored)
2hx.pers120 | External two-hour: persistence 120 min MAE | 31.675 | mg/dL | twoh_ext | artifacts/research/gt-2h-ext-2026-10-20/two_hour_external_summary.json | /persistence_mae/120 | 2026-10-20T10:10:25+00:00 | 58a1ba2570f9fe0026d918d914a11fc508b5083d03ce357bcecfe500220633ee | protocol configs/experiments/gt-2h-ext-2026-10-20.yaml (first commit c55789a 2026-10-20T15:38:49+05:30) | local / reproducible-only (git-ignored)
2hx.cmp.mean_diff | External two-hour: minus persistence at 120 min, mean_diff | -2.122 | mg/dL | twoh_ext | artifacts/research/gt-2h-ext-2026-10-20/two_hour_external_summary.json | /minus_persistence/120/mean_diff | 2026-10-20T10:10:25+00:00 | 58a1ba2570f9fe0026d918d914a11fc508b5083d03ce357bcecfe500220633ee | protocol configs/experiments/gt-2h-ext-2026-10-20.yaml (first commit c55789a 2026-10-20T15:38:49+05:30) | local / reproducible-only (git-ignored)
2hx.cmp.lo | External two-hour: minus persistence at 120 min, ci95[0] | -2.672 | mg/dL | twoh_ext | artifacts/research/gt-2h-ext-2026-10-20/two_hour_external_summary.json | /minus_persistence/120/ci95[0] | 2026-10-20T10:10:25+00:00 | 58a1ba2570f9fe0026d918d914a11fc508b5083d03ce357bcecfe500220633ee | protocol configs/experiments/gt-2h-ext-2026-10-20.yaml (first commit c55789a 2026-10-20T15:38:49+05:30) | local / reproducible-only (git-ignored)
2hx.cmp.hi | External two-hour: minus persistence at 120 min, ci95[1] | -1.526 | mg/dL | twoh_ext | artifacts/research/gt-2h-ext-2026-10-20/two_hour_external_summary.json | /minus_persistence/120/ci95[1] | 2026-10-20T10:10:25+00:00 | 58a1ba2570f9fe0026d918d914a11fc508b5083d03ce357bcecfe500220633ee | protocol configs/experiments/gt-2h-ext-2026-10-20.yaml (first commit c55789a 2026-10-20T15:38:49+05:30) | local / reproducible-only (git-ignored)
2hx.cmp.improved | External two-hour: minus persistence at 120 min, improved | 78 | people | twoh_ext | artifacts/research/gt-2h-ext-2026-10-20/two_hour_external_summary.json | /minus_persistence/120/improved | 2026-10-20T10:10:25+00:00 | 58a1ba2570f9fe0026d918d914a11fc508b5083d03ce357bcecfe500220633ee | protocol configs/experiments/gt-2h-ext-2026-10-20.yaml (first commit c55789a 2026-10-20T15:38:49+05:30) | local / reproducible-only (git-ignored)
2hx.cmp.n | External two-hour: minus persistence at 120 min, n | 98 | people | twoh_ext | artifacts/research/gt-2h-ext-2026-10-20/two_hour_external_summary.json | /minus_persistence/120/n | 2026-10-20T10:10:25+00:00 | 58a1ba2570f9fe0026d918d914a11fc508b5083d03ce357bcecfe500220633ee | protocol configs/experiments/gt-2h-ext-2026-10-20.yaml (first commit c55789a 2026-10-20T15:38:49+05:30) | local / reproducible-only (git-ignored)
2hx.cov120 | External two-hour: 80 % range coverage at 120 min | 0.87136 | share | twoh_ext | artifacts/research/gt-2h-ext-2026-10-20/two_hour_external_summary.json | /coverage/120 | 2026-10-20T10:10:25+00:00 | 58a1ba2570f9fe0026d918d914a11fc508b5083d03ce357bcecfe500220633ee | protocol configs/experiments/gt-2h-ext-2026-10-20.yaml (first commit c55789a 2026-10-20T15:38:49+05:30) | local / reproducible-only (git-ignored)
2hx.participants | External two-hour participants | 98 | people | twoh_ext | artifacts/research/gt-2h-ext-2026-10-20/two_hour_external_summary.json | /participants | 2026-10-20T10:10:25+00:00 | 58a1ba2570f9fe0026d918d914a11fc508b5083d03ce357bcecfe500220633ee | protocol configs/experiments/gt-2h-ext-2026-10-20.yaml (first commit c55789a 2026-10-20T15:38:49+05:30) | local / reproducible-only (git-ignored)
2hx.opened_at | External two-hour assessment opened | 2026-10-20 10:10 | UTC | twoh_ext | artifacts/research/gt-2h-ext-2026-10-20/two_hour_external_summary.json | /opened/opened_at | 2026-10-20T10:10:25+00:00 | 58a1ba2570f9fe0026d918d914a11fc508b5083d03ce357bcecfe500220633ee | protocol configs/experiments/gt-2h-ext-2026-10-20.yaml (first commit c55789a 2026-10-20T15:38:49+05:30) | local / reproducible-only (git-ignored)
clf.detected_onsets | Two-hour classifier (D-044): detected_onsets | 329 | onsets | clf | artifacts/research/gt-2h-clf-2026-10-20/summary.json | /results/selected/detected_onsets | 2026-10-20T09:57:29+00:00 | 6953c8edf48886d49461beae2929cd111af4eea44e40e03615a383d9f5e81a52 | protocol configs/experiments/gt-2h-clf-2026-10-20.yaml (first commit 593bdd1 2026-10-20T15:27:02+05:30) | local / reproducible-only (git-ignored)
clf.true_onsets | Two-hour classifier (D-044): true_onsets | 329 | onsets | clf | artifacts/research/gt-2h-clf-2026-10-20/summary.json | /results/selected/true_onsets | 2026-10-20T09:57:29+00:00 | 6953c8edf48886d49461beae2929cd111af4eea44e40e03615a383d9f5e81a52 | protocol configs/experiments/gt-2h-clf-2026-10-20.yaml (first commit 593bdd1 2026-10-20T15:27:02+05:30) | local / reproducible-only (git-ignored)
clf.false_alarms_per_observed_day | Two-hour classifier (D-044): false_alarms_per_observed_day | 0.181 | per observed day | clf | artifacts/research/gt-2h-clf-2026-10-20/summary.json | /results/selected/false_alarms_per_observed_day | 2026-10-20T09:57:29+00:00 | 6953c8edf48886d49461beae2929cd111af4eea44e40e03615a383d9f5e81a52 | protocol configs/experiments/gt-2h-clf-2026-10-20.yaml (first commit 593bdd1 2026-10-20T15:27:02+05:30) | local / reproducible-only (git-ignored)
clf.share_of_eligible_time_warning | Two-hour classifier (D-044): share_of_eligible_time_warning | 0.99 | share | clf | artifacts/research/gt-2h-clf-2026-10-20/summary.json | /results/selected/share_of_eligible_time_warning | 2026-10-20T09:57:29+00:00 | 6953c8edf48886d49461beae2929cd111af4eea44e40e03615a383d9f5e81a52 | protocol configs/experiments/gt-2h-clf-2026-10-20.yaml (first commit 593bdd1 2026-10-20T15:27:02+05:30) | local / reproducible-only (git-ignored)
clf.origin_roc_auc | Two-hour classifier (D-044): origin_roc_auc | 0.638 | AUC | clf | artifacts/research/gt-2h-clf-2026-10-20/summary.json | /results/selected/origin_roc_auc | 2026-10-20T09:57:29+00:00 | 6953c8edf48886d49461beae2929cd111af4eea44e40e03615a383d9f5e81a52 | protocol configs/experiments/gt-2h-clf-2026-10-20.yaml (first commit 593bdd1 2026-10-20T15:27:02+05:30) | local / reproducible-only (git-ignored)
clf.opened_at | Classifier assessment opened | 2026-10-20 09:57 | UTC | clf | artifacts/research/gt-2h-clf-2026-10-20/summary.json | /opened/opened_at | 2026-10-20T09:57:29+00:00 | 6953c8edf48886d49461beae2929cd111af4eea44e40e03615a383d9f5e81a52 | protocol configs/experiments/gt-2h-clf-2026-10-20.yaml (first commit 593bdd1 2026-10-20T15:27:02+05:30) | local / reproducible-only (git-ignored)
clf.reopened_at | Classifier assessment re-opened (post-hoc diagnostic) | 2026-10-20 09:58 | UTC | clf | artifacts/research/gt-2h-clf-2026-10-20/summary.json | /opened/reopened[0]/at | 2026-10-20T09:57:29+00:00 | 6953c8edf48886d49461beae2929cd111af4eea44e40e03615a383d9f5e81a52 | protocol configs/experiments/gt-2h-clf-2026-10-20.yaml (first commit 593bdd1 2026-10-20T15:27:02+05:30) | local / reproducible-only (git-ignored)
wr.b0.cov60 | D-046 ranges B0_calibration, 60 min coverage | 0.78237 | share | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /ranges/B0_calibration/60/coverage_pooled | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.b0.width60 | D-046 ranges B0_calibration, 60 min mean width | 66.4053 | mg/dL | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /ranges/B0_calibration/60/width_mean | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.b0.cov120 | D-046 ranges B0_calibration, 120 min coverage | 0.76488 | share | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /ranges/B0_calibration/120/coverage_pooled | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.b0.width120 | D-046 ranges B0_calibration, 120 min mean width | 88.451 | mg/dL | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /ranges/B0_calibration/120/width_mean | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.b1.cov60 | D-046 ranges B1_pooled, 60 min coverage | 0.79549 | share | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /ranges/B1_pooled/60/coverage_pooled | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.b1.width60 | D-046 ranges B1_pooled, 60 min mean width | 68.4511 | mg/dL | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /ranges/B1_pooled/60/width_mean | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.b1.cov120 | D-046 ranges B1_pooled, 120 min coverage | 0.79812 | share | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /ranges/B1_pooled/120/coverage_pooled | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.b1.width120 | D-046 ranges B1_pooled, 120 min mean width | 95.1083 | mg/dL | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /ranges/B1_pooled/120/width_mean | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.b1.pmed120 | D-046 pooled ranges: per-person median coverage, 120 min | 0.81162 | share | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /ranges/B1_pooled/120/coverage_participant_median | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.early.sel | D-046 A: onsets warned ≥ 30 min ahead, selected arm | 126 | onsets | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /warnings/selected/early | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.early.inc | D-046 A: onsets warned ≥ 30 min ahead, incumbent | 123 | onsets | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /warnings/incumbent/early | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.onsets | D-046 A: onsets (denominator) | 329 | onsets | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /warnings/selected/onsets | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.lead.sel | D-046 A: median lead, selected arm | 15 | min | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /warnings/selected/lead_median | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.diff.d | D-046 A: early-share difference selected − incumbent | 0.0091185 | share | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /warning_diff_selected_minus_incumbent/early_share_diff | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.diff.lo | 95 % CI lower | -0.054849 | share | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /warning_diff_selected_minus_incumbent/early_share_ci95[0] | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.diff.hi | 95 % CI upper | 0.066322 | share | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /warning_diff_selected_minus_incumbent/early_share_ci95[1] | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.prior.all.d | D-046 C: profile prior minus population, 120 min (all) | -0.53206 | mg/dL | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /profile_prior/all/120/c1_minus_population/mean_diff | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.prior.all.lo | 95 % CI lower | -1.572 | mg/dL | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /profile_prior/all/120/c1_minus_population/ci95[0] | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.prior.all.hi | 95 % CI upper | 0.5826 | mg/dL | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /profile_prior/all/120/c1_minus_population/ci95[1] | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.prior.d1.d | D-046 C: profile prior minus population, 120 min (first_24h) | -0.92751 | mg/dL | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /profile_prior/first_24h/120/c1_minus_population/mean_diff | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.prior.d1.lo | 95 % CI lower | -2.7386 | mg/dL | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /profile_prior/first_24h/120/c1_minus_population/ci95[0] | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.prior.d1.hi | 95 % CI upper | 0.48115 | mg/dL | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /profile_prior/first_24h/120/c1_minus_population/ci95[1] | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
wr.opened_at | D-046 outer predictions read | 2026-10-20 12:24 | UTC | warnrange | artifacts/research/gt-warn-range-2026-10-20/summary.json | /opened/opened_at | 2026-10-20T12:24:25+00:00 | cd3ea85aaca21f3765a2e9b12d0be3118c1c19e7a788b6d5d6809ccf03c421bb | protocol configs/experiments/gt-warn-range-2026-10-20.yaml (first commit aeaa93e 2026-10-20T17:51:52+05:30) | local / reproducible-only (git-ignored)
blend.improvement_60 | D-048 development check: improvement_60 (improvement positive = gain) | 0.018639 | mg/dL | blend | artifacts/research/gt-blend-2026-10-21/development.json | /decision/checks/improvement_60 | 2026-10-20T15:06:54+00:00 (file time) | f422cff1c14118d0134d4de1a2a42d30a3e843cf4b2a46da1e7faa68bb4593a4 | protocol configs/experiments/gt-blend-2026-10-21.yaml (first commit 2799ab3 2026-10-20T20:36:41+05:30) | local / reproducible-only (git-ignored)
blend.worsening_30 | D-048 development check: worsening_30 (improvement positive = gain) | -0.030327 | mg/dL | blend | artifacts/research/gt-blend-2026-10-21/development.json | /decision/checks/worsening_30 | 2026-10-20T15:06:54+00:00 (file time) | f422cff1c14118d0134d4de1a2a42d30a3e843cf4b2a46da1e7faa68bb4593a4 | protocol configs/experiments/gt-blend-2026-10-21.yaml (first commit 2799ab3 2026-10-20T20:36:41+05:30) | local / reproducible-only (git-ignored)
blend.high_worsening_60 | D-048 development check: high_worsening_60 (improvement positive = gain) | -0.17831 | mg/dL | blend | artifacts/research/gt-blend-2026-10-21/development.json | /decision/checks/high_worsening_60 | 2026-10-20T15:06:54+00:00 (file time) | f422cff1c14118d0134d4de1a2a42d30a3e843cf4b2a46da1e7faa68bb4593a4 | protocol configs/experiments/gt-blend-2026-10-21.yaml (first commit 2799ab3 2026-10-20T20:36:41+05:30) | local / reproducible-only (git-ignored)
blend.availability_drop_points | D-048 development check: availability_drop_points (improvement positive = gain) | 0 | points | blend | artifacts/research/gt-blend-2026-10-21/development.json | /decision/checks/availability_drop_points | 2026-10-20T15:06:54+00:00 (file time) | f422cff1c14118d0134d4de1a2a42d30a3e843cf4b2a46da1e7faa68bb4593a4 | protocol configs/experiments/gt-blend-2026-10-21.yaml (first commit 2799ab3 2026-10-20T20:36:41+05:30) | local / reproducible-only (git-ignored)
blend.alpha | D-048 selected blend weight | 0.75 | alpha | blend | artifacts/research/gt-blend-2026-10-21/development.json | /decision/final_selected_alpha | 2026-10-20T15:06:54+00:00 (file time) | f422cff1c14118d0134d4de1a2a42d30a3e843cf4b2a46da1e7faa68bb4593a4 | protocol configs/experiments/gt-blend-2026-10-21.yaml (first commit 2799ab3 2026-10-20T20:36:41+05:30) | local / reproducible-only (git-ignored)
blend.target_met | D-048 saved outcome: development target met | FALSE | flag | blend | artifacts/research/gt-blend-2026-10-21/development.json | /decision/target_met | 2026-10-20T15:06:54+00:00 (file time) | f422cff1c14118d0134d4de1a2a42d30a3e843cf4b2a46da1e7faa68bb4593a4 | protocol configs/experiments/gt-blend-2026-10-21.yaml (first commit 2799ab3 2026-10-20T20:36:41+05:30) | local / reproducible-only (git-ignored)
blend.inc60 | D-048 alpha = 1 (incumbent) 60-min development MAE | 22.2785 | mg/dL | blend | artifacts/research/gt-blend-2026-10-21/development.json | /folds/final/alphas/1.0/mae_60 | 2026-10-20T15:06:54+00:00 (file time) | f422cff1c14118d0134d4de1a2a42d30a3e843cf4b2a46da1e7faa68bb4593a4 | protocol configs/experiments/gt-blend-2026-10-21.yaml (first commit 2799ab3 2026-10-20T20:36:41+05:30) | local / reproducible-only (git-ignored)
blend.written_at | D-048 protocol written (before any alpha < 1 scored) | 2026-10-20 15:06 | UTC | blend_proto | artifacts/research/gt-blend-2026-10-21/protocol_written.json | /written_at | 2026-10-20T15:06:25+00:00 | acf506c74a5ab5648574e9993bf554e8ede6fbf5e4e916ebf72119e70ed68aed | protocol configs/experiments/gt-blend-2026-10-21.yaml (first commit 2799ab3 2026-10-20T20:36:41+05:30) | local / reproducible-only (git-ignored)
blend.dev_time | D-048 development decision written | 2026-10-20 15:07 | UTC | blend | artifacts/research/gt-blend-2026-10-21/development.json | file modification time of development.json (local; the file has no timestamp field) | 2026-10-20T15:06:54+00:00 (file time) | f422cff1c14118d0134d4de1a2a42d30a3e843cf4b2a46da1e7faa68bb4593a4 | protocol configs/experiments/gt-blend-2026-10-21.yaml (first commit 2799ab3 2026-10-20T20:36:41+05:30) | local / reproducible-only (git-ignored)
rt.A1.gain_60 | D-049 A1_participant_balanced vs B0: gain_60 (gain positive = lower error) | -0.039483 | mg/dL | retrain | artifacts/research/gt-retrain-2026-10-21/development.json | /selection/by_arm/A1_participant_balanced/checks/gain_60 | 2026-10-20T17:02:49+00:00 | 5531fea058f80e4638d34b485038e9a5db1b0998f3754bb24a7e39a0d63dc859 | protocol configs/experiments/gt-retrain-2026-10-21.yaml (first commit 83b36ec 2026-10-20T22:30:08+05:30) | local / reproducible-only (git-ignored)
rt.A1.worsening_30 | D-049 A1_participant_balanced vs B0: worsening_30 (gain positive = lower error) | 0.0071327 | mg/dL | retrain | artifacts/research/gt-retrain-2026-10-21/development.json | /selection/by_arm/A1_participant_balanced/checks/worsening_30 | 2026-10-20T17:02:49+00:00 | 5531fea058f80e4638d34b485038e9a5db1b0998f3754bb24a7e39a0d63dc859 | protocol configs/experiments/gt-retrain-2026-10-21.yaml (first commit 83b36ec 2026-10-20T22:30:08+05:30) | local / reproducible-only (git-ignored)
rt.A1.high_worsening_60 | D-049 A1_participant_balanced vs B0: high_worsening_60 (gain positive = lower error) | 0.0032799 | mg/dL | retrain | artifacts/research/gt-retrain-2026-10-21/development.json | /selection/by_arm/A1_participant_balanced/checks/high_worsening_60 | 2026-10-20T17:02:49+00:00 | 5531fea058f80e4638d34b485038e9a5db1b0998f3754bb24a7e39a0d63dc859 | protocol configs/experiments/gt-retrain-2026-10-21.yaml (first commit 83b36ec 2026-10-20T22:30:08+05:30) | local / reproducible-only (git-ignored)
rt.A1.availability_drop_points | D-049 A1_participant_balanced vs B0: availability_drop_points (gain positive = lower error) | 0 | points | retrain | artifacts/research/gt-retrain-2026-10-21/development.json | /selection/by_arm/A1_participant_balanced/checks/availability_drop_points | 2026-10-20T17:02:49+00:00 | 5531fea058f80e4638d34b485038e9a5db1b0998f3754bb24a7e39a0d63dc859 | protocol configs/experiments/gt-retrain-2026-10-21.yaml (first commit 83b36ec 2026-10-20T22:30:08+05:30) | local / reproducible-only (git-ignored)
rt.A1.qualifies | D-049 A1_participant_balanced: saved qualifies flag | FALSE | flag | retrain | artifacts/research/gt-retrain-2026-10-21/development.json | /selection/by_arm/A1_participant_balanced/qualifies | 2026-10-20T17:02:49+00:00 | 5531fea058f80e4638d34b485038e9a5db1b0998f3754bb24a7e39a0d63dc859 | protocol configs/experiments/gt-retrain-2026-10-21.yaml (first commit 83b36ec 2026-10-20T22:30:08+05:30) | local / reproducible-only (git-ignored)
rt.A2.gain_60 | D-049 A2_balanced_rising_2x vs B0: gain_60 (gain positive = lower error) | 0.070008 | mg/dL | retrain | artifacts/research/gt-retrain-2026-10-21/development.json | /selection/by_arm/A2_balanced_rising_2x/checks/gain_60 | 2026-10-20T17:02:49+00:00 | 5531fea058f80e4638d34b485038e9a5db1b0998f3754bb24a7e39a0d63dc859 | protocol configs/experiments/gt-retrain-2026-10-21.yaml (first commit 83b36ec 2026-10-20T22:30:08+05:30) | local / reproducible-only (git-ignored)
rt.A2.worsening_30 | D-049 A2_balanced_rising_2x vs B0: worsening_30 (gain positive = lower error) | 0.047787 | mg/dL | retrain | artifacts/research/gt-retrain-2026-10-21/development.json | /selection/by_arm/A2_balanced_rising_2x/checks/worsening_30 | 2026-10-20T17:02:49+00:00 | 5531fea058f80e4638d34b485038e9a5db1b0998f3754bb24a7e39a0d63dc859 | protocol configs/experiments/gt-retrain-2026-10-21.yaml (first commit 83b36ec 2026-10-20T22:30:08+05:30) | local / reproducible-only (git-ignored)
rt.A2.high_worsening_60 | D-049 A2_balanced_rising_2x vs B0: high_worsening_60 (gain positive = lower error) | 0.16537 | mg/dL | retrain | artifacts/research/gt-retrain-2026-10-21/development.json | /selection/by_arm/A2_balanced_rising_2x/checks/high_worsening_60 | 2026-10-20T17:02:49+00:00 | 5531fea058f80e4638d34b485038e9a5db1b0998f3754bb24a7e39a0d63dc859 | protocol configs/experiments/gt-retrain-2026-10-21.yaml (first commit 83b36ec 2026-10-20T22:30:08+05:30) | local / reproducible-only (git-ignored)
rt.A2.availability_drop_points | D-049 A2_balanced_rising_2x vs B0: availability_drop_points (gain positive = lower error) | 0 | points | retrain | artifacts/research/gt-retrain-2026-10-21/development.json | /selection/by_arm/A2_balanced_rising_2x/checks/availability_drop_points | 2026-10-20T17:02:49+00:00 | 5531fea058f80e4638d34b485038e9a5db1b0998f3754bb24a7e39a0d63dc859 | protocol configs/experiments/gt-retrain-2026-10-21.yaml (first commit 83b36ec 2026-10-20T22:30:08+05:30) | local / reproducible-only (git-ignored)
rt.A2.qualifies | D-049 A2_balanced_rising_2x: saved qualifies flag | FALSE | flag | retrain | artifacts/research/gt-retrain-2026-10-21/development.json | /selection/by_arm/A2_balanced_rising_2x/qualifies | 2026-10-20T17:02:49+00:00 | 5531fea058f80e4638d34b485038e9a5db1b0998f3754bb24a7e39a0d63dc859 | protocol configs/experiments/gt-retrain-2026-10-21.yaml (first commit 83b36ec 2026-10-20T22:30:08+05:30) | local / reproducible-only (git-ignored)
rt.B1.gain_60 | D-049 B1_balanced_long_history vs B0: gain_60 (gain positive = lower error) | 0.41122 | mg/dL | retrain | artifacts/research/gt-retrain-2026-10-21/development.json | /selection/by_arm/B1_balanced_long_history/checks/gain_60 | 2026-10-20T17:02:49+00:00 | 5531fea058f80e4638d34b485038e9a5db1b0998f3754bb24a7e39a0d63dc859 | protocol configs/experiments/gt-retrain-2026-10-21.yaml (first commit 83b36ec 2026-10-20T22:30:08+05:30) | local / reproducible-only (git-ignored)
rt.B1.worsening_30 | D-049 B1_balanced_long_history vs B0: worsening_30 (gain positive = lower error) | -0.085897 | mg/dL | retrain | artifacts/research/gt-retrain-2026-10-21/development.json | /selection/by_arm/B1_balanced_long_history/checks/worsening_30 | 2026-10-20T17:02:49+00:00 | 5531fea058f80e4638d34b485038e9a5db1b0998f3754bb24a7e39a0d63dc859 | protocol configs/experiments/gt-retrain-2026-10-21.yaml (first commit 83b36ec 2026-10-20T22:30:08+05:30) | local / reproducible-only (git-ignored)
rt.B1.high_worsening_60 | D-049 B1_balanced_long_history vs B0: high_worsening_60 (gain positive = lower error) | -0.18038 | mg/dL | retrain | artifacts/research/gt-retrain-2026-10-21/development.json | /selection/by_arm/B1_balanced_long_history/checks/high_worsening_60 | 2026-10-20T17:02:49+00:00 | 5531fea058f80e4638d34b485038e9a5db1b0998f3754bb24a7e39a0d63dc859 | protocol configs/experiments/gt-retrain-2026-10-21.yaml (first commit 83b36ec 2026-10-20T22:30:08+05:30) | local / reproducible-only (git-ignored)
rt.B1.availability_drop_points | D-049 B1_balanced_long_history vs B0: availability_drop_points (gain positive = lower error) | 0 | points | retrain | artifacts/research/gt-retrain-2026-10-21/development.json | /selection/by_arm/B1_balanced_long_history/checks/availability_drop_points | 2026-10-20T17:02:49+00:00 | 5531fea058f80e4638d34b485038e9a5db1b0998f3754bb24a7e39a0d63dc859 | protocol configs/experiments/gt-retrain-2026-10-21.yaml (first commit 83b36ec 2026-10-20T22:30:08+05:30) | local / reproducible-only (git-ignored)
rt.B1.qualifies | D-049 B1_balanced_long_history: saved qualifies flag | FALSE | flag | retrain | artifacts/research/gt-retrain-2026-10-21/development.json | /selection/by_arm/B1_balanced_long_history/qualifies | 2026-10-20T17:02:49+00:00 | 5531fea058f80e4638d34b485038e9a5db1b0998f3754bb24a7e39a0d63dc859 | protocol configs/experiments/gt-retrain-2026-10-21.yaml (first commit 83b36ec 2026-10-20T22:30:08+05:30) | local / reproducible-only (git-ignored)
rt.B1pop.d | D-049 B1 population-only 60-min change (descriptive, after selection) | -1.3872 | mg/dL | retrain | artifacts/research/gt-retrain-2026-10-21/development.json | /development_paired_descriptive/B1_pop_60/mean_diff | 2026-10-20T17:02:49+00:00 | 5531fea058f80e4638d34b485038e9a5db1b0998f3754bb24a7e39a0d63dc859 | protocol configs/experiments/gt-retrain-2026-10-21.yaml (first commit 83b36ec 2026-10-20T22:30:08+05:30) | local / reproducible-only (git-ignored)
rt.B1pop.lo | 95 % CI lower | -2.6219 | mg/dL | retrain | artifacts/research/gt-retrain-2026-10-21/development.json | /development_paired_descriptive/B1_pop_60/ci95[0] | 2026-10-20T17:02:49+00:00 | 5531fea058f80e4638d34b485038e9a5db1b0998f3754bb24a7e39a0d63dc859 | protocol configs/experiments/gt-retrain-2026-10-21.yaml (first commit 83b36ec 2026-10-20T22:30:08+05:30) | local / reproducible-only (git-ignored)
rt.B1pop.hi | 95 % CI upper | -0.29174 | mg/dL | retrain | artifacts/research/gt-retrain-2026-10-21/development.json | /development_paired_descriptive/B1_pop_60/ci95[1] | 2026-10-20T17:02:49+00:00 | 5531fea058f80e4638d34b485038e9a5db1b0998f3754bb24a7e39a0d63dc859 | protocol configs/experiments/gt-retrain-2026-10-21.yaml (first commit 83b36ec 2026-10-20T22:30:08+05:30) | local / reproducible-only (git-ignored)
rt.started_at | D-049 run start | 2026-10-20 17:03 | UTC | retrain | artifacts/research/gt-retrain-2026-10-21/development.json | /started_at | 2026-10-20T17:02:49+00:00 | 5531fea058f80e4638d34b485038e9a5db1b0998f3754bb24a7e39a0d63dc859 | protocol configs/experiments/gt-retrain-2026-10-21.yaml (first commit 83b36ec 2026-10-20T22:30:08+05:30) | local / reproducible-only (git-ignored)
rg.err.G | D-050 pre-test mean ¦coverage − 0.80¦, shape G | 0.064843 | share | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /pre_test_check/coverage_error/G | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.width.G | D-050 pre-test mean width, shape G | 54.9728 | mg/dL | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /pre_test_check/width_mean/G | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.err.T | D-050 pre-test mean ¦coverage − 0.80¦, shape T | 0.057674 | share | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /pre_test_check/coverage_error/T | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.width.T | D-050 pre-test mean width, shape T | 51.5946 | mg/dL | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /pre_test_check/width_mean/T | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.chosen | D-050 saved choice | G | text | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /pre_test_check/chosen | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.inc.cov30 | D-050 final test (read once), incumbent_G_calibration, 30 min coverage | 0.74204 | share | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/incumbent_G_calibration/30/coverage | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.inc.width30 | D-050 final test, incumbent_G_calibration, 30 min mean width | 33.4234 | mg/dL | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/incumbent_G_calibration/30/width_mean | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.inc.rising30 | D-050 final test, incumbent_G_calibration, 30 min, rising (> 1 mg/dL/min) coverage | 0.4556 | share | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/incumbent_G_calibration/30/by_stratum/rising/coverage | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.inc.risingn30 | D-050 final test, rising, 30 min origins | 687 | origins | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/incumbent_G_calibration/30/by_stratum/rising/n | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.inc.steady30 | D-050 final test, incumbent_G_calibration, 30 min, steady (> 1 mg/dL/min) coverage | 0.78394 | share | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/incumbent_G_calibration/30/by_stratum/steady/coverage | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.inc.steadyn30 | D-050 final test, steady, 30 min origins | 6350 | origins | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/incumbent_G_calibration/30/by_stratum/steady/n | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.inc.falling30 | D-050 final test, incumbent_G_calibration, 30 min, falling (> 1 mg/dL/min) coverage | 0.6383 | share | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/incumbent_G_calibration/30/by_stratum/falling/coverage | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.inc.fallingn30 | D-050 final test, falling, 30 min origins | 658 | origins | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/incumbent_G_calibration/30/by_stratum/falling/n | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.inc.cov60 | D-050 final test (read once), incumbent_G_calibration, 60 min coverage | 0.73979 | share | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/incumbent_G_calibration/60/coverage | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.inc.width60 | D-050 final test, incumbent_G_calibration, 60 min mean width | 52.7294 | mg/dL | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/incumbent_G_calibration/60/width_mean | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.inc.rising60 | D-050 final test, incumbent_G_calibration, 60 min, rising (> 1 mg/dL/min) coverage | 0.51622 | share | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/incumbent_G_calibration/60/by_stratum/rising/coverage | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.inc.risingn60 | D-050 final test, rising, 60 min origins | 678 | origins | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/incumbent_G_calibration/60/by_stratum/rising/n | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.inc.steady60 | D-050 final test, incumbent_G_calibration, 60 min, steady (> 1 mg/dL/min) coverage | 0.77942 | share | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/incumbent_G_calibration/60/by_stratum/steady/coverage | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.inc.steadyn60 | D-050 final test, steady, 60 min origins | 6347 | origins | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/incumbent_G_calibration/60/by_stratum/steady/n | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.inc.falling60 | D-050 final test, incumbent_G_calibration, 60 min, falling (> 1 mg/dL/min) coverage | 0.59002 | share | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/incumbent_G_calibration/60/by_stratum/falling/coverage | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.inc.fallingn60 | D-050 final test, falling, 60 min origins | 661 | origins | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/incumbent_G_calibration/60/by_stratum/falling/n | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.pool.cov30 | D-050 final test (read once), G_pooled, 30 min coverage | 0.77919 | share | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/G_pooled/30/coverage | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.pool.width30 | D-050 final test, G_pooled, 30 min mean width | 37.7691 | mg/dL | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/G_pooled/30/width_mean | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.pool.rising30 | D-050 final test, G_pooled, 30 min, rising (> 1 mg/dL/min) coverage | 0.49636 | share | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/G_pooled/30/by_stratum/rising/coverage | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.pool.steady30 | D-050 final test, G_pooled, 30 min, steady (> 1 mg/dL/min) coverage | 0.8178 | share | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/G_pooled/30/by_stratum/steady/coverage | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.pool.falling30 | D-050 final test, G_pooled, 30 min, falling (> 1 mg/dL/min) coverage | 0.70365 | share | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/G_pooled/30/by_stratum/falling/coverage | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.pool.cov60 | D-050 final test (read once), G_pooled, 60 min coverage | 0.78687 | share | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/G_pooled/60/coverage | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.pool.width60 | D-050 final test, G_pooled, 60 min mean width | 59.5413 | mg/dL | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/G_pooled/60/width_mean | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.pool.rising60 | D-050 final test, G_pooled, 60 min, rising (> 1 mg/dL/min) coverage | 0.5649 | share | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/G_pooled/60/by_stratum/rising/coverage | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.pool.steady60 | D-050 final test, G_pooled, 60 min, steady (> 1 mg/dL/min) coverage | 0.8248 | share | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/G_pooled/60/by_stratum/steady/coverage | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.pool.falling60 | D-050 final test, G_pooled, 60 min, falling (> 1 mg/dL/min) coverage | 0.65204 | share | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /test_retrospective/G_pooled/60/by_stratum/falling/coverage | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
rg.generated_at | D-050 run time | 2026-10-20 17:40 | UTC | ranges | artifacts/research/gt-ranges-2026-10-21/summary.json | /generated_at | 2026-10-20T17:40:17+00:00 | 181b8c28e242361beeda778b2f9baa4b7a1091cf66e2e14b7a178682de8df7fd | protocol configs/experiments/gt-ranges-2026-10-21.yaml (first commit 0de0889 2026-10-20T23:09:42+05:30) | local / reproducible-only (git-ignored)
tu.C0.mae60 | D-051 C0_current_features: 60-min personalised macro MAE | 22.2668 | mg/dL | tune | artifacts/research/gt-tune-2026-10-21/summary.json | /arms/C0_current_features/60/mae | 2026-10-20T17:46:24+00:00 | fa860b4c52b4ddbf583de4bfc3ce57e24ca2a10aaee5aaffbd25c71ff2969bd7 | protocol configs/experiments/gt-tune-2026-10-21.yaml (first commit e7ff594 2026-10-20T23:14:37+05:30) | local / reproducible-only (git-ignored)
tu.C1.mae60 | D-051 C1_with_history: 60-min personalised macro MAE | 21.9006 | mg/dL | tune | artifacts/research/gt-tune-2026-10-21/summary.json | /arms/C1_with_history/60/mae | 2026-10-20T17:46:24+00:00 | fa860b4c52b4ddbf583de4bfc3ce57e24ca2a10aaee5aaffbd25c71ff2969bd7 | protocol configs/experiments/gt-tune-2026-10-21.yaml (first commit e7ff594 2026-10-20T23:14:37+05:30) | local / reproducible-only (git-ignored)
tu.R0.mae60 | D-051 R0_untuned_serving_recipe: 60-min personalised macro MAE | 22.5482 | mg/dL | tune | artifacts/research/gt-tune-2026-10-21/summary.json | /arms/R0_untuned_serving_recipe/60/mae | 2026-10-20T17:46:24+00:00 | fa860b4c52b4ddbf583de4bfc3ce57e24ca2a10aaee5aaffbd25c71ff2969bd7 | protocol configs/experiments/gt-tune-2026-10-21.yaml (first commit e7ff594 2026-10-20T23:14:37+05:30) | local / reproducible-only (git-ignored)
tu.P.mae60 | D-051 persistence: 60-min personalised macro MAE | 24.4672 | mg/dL | tune | artifacts/research/gt-tune-2026-10-21/summary.json | /arms/persistence/60/mae | 2026-10-20T17:46:24+00:00 | fa860b4c52b4ddbf583de4bfc3ce57e24ca2a10aaee5aaffbd25c71ff2969bd7 | protocol configs/experiments/gt-tune-2026-10-21.yaml (first commit e7ff594 2026-10-20T23:14:37+05:30) | local / reproducible-only (git-ignored)
tu.c1c0.mean_diff | D-051 paired C1_with_history¦C0_current_features¦final¦60: mean_diff | -0.36626 | mg/dL | tune | artifacts/research/gt-tune-2026-10-21/summary.json | /paired/C1_with_history¦C0_current_features¦final¦60/mean_diff | 2026-10-20T17:46:24+00:00 | fa860b4c52b4ddbf583de4bfc3ce57e24ca2a10aaee5aaffbd25c71ff2969bd7 | protocol configs/experiments/gt-tune-2026-10-21.yaml (first commit e7ff594 2026-10-20T23:14:37+05:30) | local / reproducible-only (git-ignored)
tu.c1c0.lo | D-051 paired C1_with_history¦C0_current_features¦final¦60: ci95[0] | -0.6333 | mg/dL | tune | artifacts/research/gt-tune-2026-10-21/summary.json | /paired/C1_with_history¦C0_current_features¦final¦60/ci95[0] | 2026-10-20T17:46:24+00:00 | fa860b4c52b4ddbf583de4bfc3ce57e24ca2a10aaee5aaffbd25c71ff2969bd7 | protocol configs/experiments/gt-tune-2026-10-21.yaml (first commit e7ff594 2026-10-20T23:14:37+05:30) | local / reproducible-only (git-ignored)
tu.c1c0.hi | D-051 paired C1_with_history¦C0_current_features¦final¦60: ci95[1] | -0.05907 | mg/dL | tune | artifacts/research/gt-tune-2026-10-21/summary.json | /paired/C1_with_history¦C0_current_features¦final¦60/ci95[1] | 2026-10-20T17:46:24+00:00 | fa860b4c52b4ddbf583de4bfc3ce57e24ca2a10aaee5aaffbd25c71ff2969bd7 | protocol configs/experiments/gt-tune-2026-10-21.yaml (first commit e7ff594 2026-10-20T23:14:37+05:30) | local / reproducible-only (git-ignored)
tu.c1c0.improved | D-051 paired C1_with_history¦C0_current_features¦final¦60: improved | 10 | people | tune | artifacts/research/gt-tune-2026-10-21/summary.json | /paired/C1_with_history¦C0_current_features¦final¦60/improved | 2026-10-20T17:46:24+00:00 | fa860b4c52b4ddbf583de4bfc3ce57e24ca2a10aaee5aaffbd25c71ff2969bd7 | protocol configs/experiments/gt-tune-2026-10-21.yaml (first commit e7ff594 2026-10-20T23:14:37+05:30) | local / reproducible-only (git-ignored)
tu.c1r0.mean_diff | D-051 paired C1_with_history¦R0_untuned_serving_recipe¦final¦60: mean_diff | -0.64765 | mg/dL | tune | artifacts/research/gt-tune-2026-10-21/summary.json | /paired/C1_with_history¦R0_untuned_serving_recipe¦final¦60/mean_diff | 2026-10-20T17:46:24+00:00 | fa860b4c52b4ddbf583de4bfc3ce57e24ca2a10aaee5aaffbd25c71ff2969bd7 | protocol configs/experiments/gt-tune-2026-10-21.yaml (first commit e7ff594 2026-10-20T23:14:37+05:30) | local / reproducible-only (git-ignored)
tu.c1r0.lo | D-051 paired C1_with_history¦R0_untuned_serving_recipe¦final¦60: ci95[0] | -1.4072 | mg/dL | tune | artifacts/research/gt-tune-2026-10-21/summary.json | /paired/C1_with_history¦R0_untuned_serving_recipe¦final¦60/ci95[0] | 2026-10-20T17:46:24+00:00 | fa860b4c52b4ddbf583de4bfc3ce57e24ca2a10aaee5aaffbd25c71ff2969bd7 | protocol configs/experiments/gt-tune-2026-10-21.yaml (first commit e7ff594 2026-10-20T23:14:37+05:30) | local / reproducible-only (git-ignored)
tu.c1r0.hi | D-051 paired C1_with_history¦R0_untuned_serving_recipe¦final¦60: ci95[1] | 0.069974 | mg/dL | tune | artifacts/research/gt-tune-2026-10-21/summary.json | /paired/C1_with_history¦R0_untuned_serving_recipe¦final¦60/ci95[1] | 2026-10-20T17:46:24+00:00 | fa860b4c52b4ddbf583de4bfc3ce57e24ca2a10aaee5aaffbd25c71ff2969bd7 | protocol configs/experiments/gt-tune-2026-10-21.yaml (first commit e7ff594 2026-10-20T23:14:37+05:30) | local / reproducible-only (git-ignored)
tu.c1r0.improved | D-051 paired C1_with_history¦R0_untuned_serving_recipe¦final¦60: improved | 8 | people | tune | artifacts/research/gt-tune-2026-10-21/summary.json | /paired/C1_with_history¦R0_untuned_serving_recipe¦final¦60/improved | 2026-10-20T17:46:24+00:00 | fa860b4c52b4ddbf583de4bfc3ce57e24ca2a10aaee5aaffbd25c71ff2969bd7 | protocol configs/experiments/gt-tune-2026-10-21.yaml (first commit e7ff594 2026-10-20T23:14:37+05:30) | local / reproducible-only (git-ignored)
tu.adopt.worsening_30 | D-051 adoption check: worsening_30 | -0.064044 | mg/dL | tune | artifacts/research/gt-tune-2026-10-21/summary.json | /adoption/worsening_30 | 2026-10-20T17:46:24+00:00 | fa860b4c52b4ddbf583de4bfc3ce57e24ca2a10aaee5aaffbd25c71ff2969bd7 | protocol configs/experiments/gt-tune-2026-10-21.yaml (first commit e7ff594 2026-10-20T23:14:37+05:30) | local / reproducible-only (git-ignored)
tu.adopt.high_worsening_60 | D-051 adoption check: high_worsening_60 | -0.79724 | mg/dL | tune | artifacts/research/gt-tune-2026-10-21/summary.json | /adoption/high_worsening_60 | 2026-10-20T17:46:24+00:00 | fa860b4c52b4ddbf583de4bfc3ce57e24ca2a10aaee5aaffbd25c71ff2969bd7 | protocol configs/experiments/gt-tune-2026-10-21.yaml (first commit e7ff594 2026-10-20T23:14:37+05:30) | local / reproducible-only (git-ignored)
tu.adopt.c1_better | D-051 adoption check: c1_better | FALSE | flag | tune | artifacts/research/gt-tune-2026-10-21/summary.json | /adoption/c1_better | 2026-10-20T17:46:24+00:00 | fa860b4c52b4ddbf583de4bfc3ce57e24ca2a10aaee5aaffbd25c71ff2969bd7 | protocol configs/experiments/gt-tune-2026-10-21.yaml (first commit e7ff594 2026-10-20T23:14:37+05:30) | local / reproducible-only (git-ignored)
tu.opened_at | D-051 outer folds opened | 2026-10-20 17:51 | UTC | tune | artifacts/research/gt-tune-2026-10-21/summary.json | /opened/opened_at | 2026-10-20T17:46:24+00:00 | fa860b4c52b4ddbf583de4bfc3ce57e24ca2a10aaee5aaffbd25c71ff2969bd7 | protocol configs/experiments/gt-tune-2026-10-21.yaml (first commit e7ff594 2026-10-20T23:14:37+05:30) | local / reproducible-only (git-ignored)
tu.started_at | D-051 run start | 2026-10-20 17:46 | UTC | tune | artifacts/research/gt-tune-2026-10-21/summary.json | /started_at | 2026-10-20T17:46:24+00:00 | fa860b4c52b4ddbf583de4bfc3ce57e24ca2a10aaee5aaffbd25c71ff2969bd7 | protocol configs/experiments/gt-tune-2026-10-21.yaml (first commit e7ff594 2026-10-20T23:14:37+05:30) | local / reproducible-only (git-ignored)
jt.S.mae60 | D-052 S: 60-min personalised macro MAE | 21.6158 | mg/dL | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /procedures/S/60/mae | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
jt.S.share30 | D-052 S: share of all onsets warned ≥ 30 min | 0.34347 | share | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /warnings/S/share_ge30 | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
jt.S.fa | D-052 S: false alarms per observed day | 1.7894 | per observed day | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /warnings/S/fa_per_observed_day | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
jt.R0.mae60 | D-052 R0: 60-min personalised macro MAE | 22.4998 | mg/dL | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /procedures/R0/60/mae | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
jt.R0.share30 | D-052 R0: share of all onsets warned ≥ 30 min | 0.18845 | share | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /warnings/R0/share_ge30 | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
jt.R0.fa | D-052 R0: false alarms per observed day | 1.1881 | per observed day | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /warnings/R0/fa_per_observed_day | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
jt.gate.mean_diff_60 | D-052 primary gate value: mean_diff_60 | -0.88406 | mg/dL | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /gate/mean_diff_60 | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
jt.gate.ci95_60_lo | D-052 primary gate value: ci95_60[0] | -1.7536 | mg/dL | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /gate/ci95_60[0] | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
jt.gate.ci95_60_hi | D-052 primary gate value: ci95_60[1] | -0.13303 | mg/dL | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /gate/ci95_60[1] | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
jt.gate.worsening_30 | D-052 primary gate value: worsening_30 | -0.50091 | mg/dL | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /gate/worsening_30 | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
jt.gate.high_worsening_60 | D-052 primary gate value: high_worsening_60 | 1.6971 | mg/dL | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /gate/high_worsening_60 | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
jt.gate.availability_drop_points | D-052 primary gate value: availability_drop_points | 0 | mg/dL | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /gate/availability_drop_points | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
jt.gate.passed | D-052 primary gate value: passed | FALSE | flag | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /gate/passed | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
jt.rel.fa_per_day_delta | D-052 release guardrail value: fa_per_day_delta | 0.6013 |  | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /release_guardrails/fa_per_day_delta | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
jt.rel.time_in_warning_delta_points | D-052 release guardrail value: time_in_warning_delta_points | 5.3572 |  | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /release_guardrails/time_in_warning_delta_points | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
jt.rel.warned30_delta_points | D-052 release guardrail value: warned30_delta_points | 15.5015 |  | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /release_guardrails/warned30_delta_points | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
jt.rel.passed | D-052 release guardrail value: passed | FALSE | flag | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /release_guardrails/passed | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
jt.cov30 | D-052 candidate S range coverage, 30 min | 0.79864 | share | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /ranges/S/30/coverage | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
jt.cov60 | D-052 candidate S range coverage, 60 min | 0.78414 | share | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /ranges/S/60/coverage | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
jt.started_at | D-052 run start | 2026-10-20 18:44 | UTC | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /started_at | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
jt.opened_at | D-052 outer folds opened | 2026-10-20 18:57 | UTC | joint | artifacts/research/gt-joint-2026-10-21/summary.json | /opened/opened_at | 2026-10-20T18:44:07+00:00 | 0feafab67e20ce254f0c6e89e6f6d6a8b3a66d0019d28d2786e03ffc75312bdb | protocol configs/experiments/gt-joint-2026-10-21.yaml (first commit 903419d 2026-10-21T00:10:51+05:30) | local / reproducible-only (git-ignored)
cs.S.mae30 | D-053 S: 30-min personalised macro MAE | 13.8917 | mg/dL | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /procedures/S/30/mae | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.S.mae60 | D-053 S: 60-min personalised macro MAE | 22.0032 | mg/dL | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /procedures/S/60/mae | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.R0.mae30 | D-053 R0: 30-min personalised macro MAE | 14.391 | mg/dL | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /procedures/R0/30/mae | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.R0.mae60 | D-053 R0: 60-min personalised macro MAE | 23.1954 | mg/dL | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /procedures/R0/60/mae | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.persistence.mae30 | D-053 persistence: 30-min personalised macro MAE | 15.3387 | mg/dL | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /procedures/persistence/30/mae | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.persistence.mae60 | D-053 persistence: 60-min personalised macro MAE | 24.4672 | mg/dL | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /procedures/persistence/60/mae | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.S.high60 | D-053 S: 60-min macro MAE on targets > 180 | 39.454 | mg/dL | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /procedures/S/60/mae_high | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.R0.high60 | D-053 R0: 60-min macro MAE on targets > 180 | 38.3882 | mg/dL | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /procedures/R0/60/mae_high | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.p60.mean_diff | D-053 paired S − R0 at 60 min: mean_diff | -1.1923 | mg/dL | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /paired/S¦R0¦final¦60/mean_diff | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.p60.lo | D-053 paired S − R0 at 60 min: ci95[0] | -2.0947 | mg/dL | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /paired/S¦R0¦final¦60/ci95[0] | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.p60.hi | D-053 paired S − R0 at 60 min: ci95[1] | -0.38742 | mg/dL | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /paired/S¦R0¦final¦60/ci95[1] | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.p60.improved | D-053 paired S − R0 at 60 min: improved | 11 | people | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /paired/S¦R0¦final¦60/improved | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.pg.mean_diff_60 | D-053 primary gate value (unrounded): mean_diff_60 | -1.1923 | mg/dL | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /primary_gate/values/mean_diff_60 | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.pgok.mean_diff_60 | D-053 saved check outcome: mean_diff_60 | TRUE | flag | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /primary_gate/checks/mean_diff_60 | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.pg.ci95_60_hi | D-053 primary gate value (unrounded): ci95_60[1] | -0.38742 | mg/dL | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /primary_gate/values/ci95_60[1] | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.pgok.ci95_upper_below_0 | D-053 saved check outcome: ci95_60[1] | TRUE | flag | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /primary_gate/checks/ci95_upper_below_0 | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.pg.worsening_30 | D-053 primary gate value (unrounded): worsening_30 | -0.49933 | mg/dL | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /primary_gate/values/worsening_30 | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.pgok.worsening_30 | D-053 saved check outcome: worsening_30 | TRUE | flag | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /primary_gate/checks/worsening_30 | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.pg.high_worsening_60 | D-053 primary gate value (unrounded): high_worsening_60 | 1.0658 | mg/dL | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /primary_gate/values/high_worsening_60 | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.pgok.high_worsening_60 | D-053 saved check outcome: high_worsening_60 | FALSE | flag | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /primary_gate/checks/high_worsening_60 | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.pg.availability_drop_points | D-053 primary gate value (unrounded): availability_drop_points | 0 | points | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /primary_gate/values/availability_drop_points | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.pgok.availability_drop | D-053 saved check outcome: availability_drop_points | TRUE | flag | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /primary_gate/checks/availability_drop | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.pg.passed | D-053 saved primary gate outcome | FALSE | flag | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /primary_gate/passed | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.w.S.fa_day | D-053 warning S: fa_day | 1.1519 | per observed day | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /warnings/S/fa_day | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.w.S.share30 | D-053 warning S: share30 | 0.21884 | share | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /warnings/S/share30 | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.w.S.tiw | D-053 warning S: tiw | 0.060702 | share | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /warnings/S/tiw | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.w.S.false_alarms | D-053 warning S: false_alarms | 159 | episodes | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /warnings/S/false_alarms | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.w.S.observed_days | D-053 warning S: observed_days | 138.0347 | patient-days | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /warnings/S/observed_days | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.w.S.onsets | D-053 warning S: onsets | 329 | onsets | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /warnings/S/onsets | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.w.R0P.fa_day | D-053 warning R0-P: fa_day | 1.1954 | per observed day | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /warnings/R0-P/fa_day | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.w.R0P.share30 | D-053 warning R0-P: share30 | 0.16413 | share | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /warnings/R0-P/share30 | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.w.R0P.tiw | D-053 warning R0-P: tiw | 0.050096 | share | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /warnings/R0-P/tiw | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.w.R0P.false_alarms | D-053 warning R0-P: false_alarms | 165 | episodes | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /warnings/R0-P/false_alarms | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.w.R0P.observed_days | D-053 warning R0-P: observed_days | 138.0347 | patient-days | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /warnings/R0-P/observed_days | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.w.R0P.onsets | D-053 warning R0-P: onsets | 329 | onsets | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /warnings/R0-P/onsets | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.op.passed | D-053 saved operational gate outcome | TRUE | flag | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /operational_gate/passed | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.rg.passed | D-053 saved range gate outcome | TRUE | flag | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /range_gate/passed | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.all_passed | D-053 saved all-gates outcome | FALSE | flag | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /all_gates_passed | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.cov30 | D-053 S exact-predictor 80 % range coverage, 30 min | 0.79428 | share | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /ranges/S/30/coverage | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.cov60 | D-053 S exact-predictor 80 % range coverage, 60 min | 0.77261 | share | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /ranges/S/60/coverage | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.started_at | D-053 run start | 2026-10-21 17:06 | UTC | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /manifest/started_at | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.opened_at | D-053 assessment opened | 2026-10-21 18:09 | UTC | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /assessment_opened/opened_at | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
cs.fits | D-053 logical fits | 1490 | fits | constrained | artifacts/research/gt-constrained-2026-10-21/summary.json | /fit_totals/fits | 2026-10-21T17:06:28+00:00 | d5385c8859b58494541d775b8846e422ad8de48c69b520365ec325f594cf7e14 | protocol configs/experiments/gt-constrained-2026-10-21.yaml (first commit 2e2b01a 2026-10-21T22:31:57+05:30) | local / reproducible-only (git-ignored)
fd.FM60 | D-054 zero-shot foundation model, 60-min macro MAE | 22.9237 | mg/dL | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/60/FM_vs_persistence/macro/FM | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.FM30 | D-054 foundation model, 30-min macro MAE | 14.5456 | mg/dL | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/30/FM_vs_persistence/macro/FM | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.P60 | D-054 persistence on identical origins, 60 min | 24.4672 | mg/dL | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/60/FM_vs_persistence/macro/persistence | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.P30 | D-054 persistence on identical origins, 30 min | 15.3387 | mg/dL | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/30/FM_vs_persistence/macro/persistence | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.R060 | D-054 refitted serving recipe (saved D-052 R0), 60 min | 22.4998 | mg/dL | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/60/FM_vs_R0saved_final/macro/R0 | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.fp.mean | D-054 FM_vs_persistence, 60 min: mean | -1.5435 | mg/dL | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/60/FM_vs_persistence/diff_macro/mean | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.fp.lo | D-054 FM_vs_persistence, 60 min: lo | -1.8079 | mg/dL | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/60/FM_vs_persistence/diff_macro/lo | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.fp.hi | D-054 FM_vs_persistence, 60 min: hi | -1.2958 | mg/dL | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/60/FM_vs_persistence/diff_macro/hi | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.fp.people_better | D-054 FM_vs_persistence, 60 min: people_better | 14 | people | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/60/FM_vs_persistence/diff_macro/people_better | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.fr.mean | D-054 FM_vs_R0saved_final, 60 min: mean | 0.42388 | mg/dL | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/60/FM_vs_R0saved_final/diff_macro/mean | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.fr.lo | D-054 FM_vs_R0saved_final, 60 min: lo | -0.70158 | mg/dL | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/60/FM_vs_R0saved_final/diff_macro/lo | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.fr.hi | D-054 FM_vs_R0saved_final, 60 min: hi | 1.4742 | mg/dL | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/60/FM_vs_R0saved_final/diff_macro/hi | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.fr.people_better | D-054 FM_vs_R0saved_final, 60 min: people_better | 4 | people | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/60/FM_vs_R0saved_final/diff_macro/people_better | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.f1.mean | D-054 F1_final_vs_R0refit_final, 60 min: mean | -0.45983 | mg/dL | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/60/F1_final_vs_R0refit_final/diff_macro/mean | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.f1.lo | D-054 F1_final_vs_R0refit_final, 60 min: lo | -0.72001 | mg/dL | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/60/F1_final_vs_R0refit_final/diff_macro/lo | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.f1.hi | D-054 F1_final_vs_R0refit_final, 60 min: hi | -0.21345 | mg/dL | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/60/F1_final_vs_R0refit_final/diff_macro/hi | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.f1.people_better | D-054 F1_final_vs_R0refit_final, 60 min: people_better | 12 | people | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/60/F1_final_vs_R0refit_final/diff_macro/people_better | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.high.FM | D-054 foundation, targets > 180, 60 min | 43.6724 | mg/dL | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/60/FM_vs_R0saved_final/high_gt180/macro/FM | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.high.R0 | D-054 R0, targets > 180, 60 min | 38.2828 | mg/dL | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/60/FM_vs_R0saved_final/high_gt180/macro/R0 | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.generated_at | D-054 summary generated | 2026-10-21 16:10 | UTC | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /generated_at | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.checkpoint | D-054 checkpoint | amazon/chronos-bolt-tiny | text | foundation_meta | artifacts/research/gt-foundation-2026-10-21/run_meta.json | /checkpoint | 2026-10-21T16:09:31+00:00 (file time) | 08755c5f2897c441bd0bd4c30577e8151fc571e64000bd6625458af4f33c3b1e | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.revision | D-054 pinned checkpoint revision | a0e552de83495b5c28c14c71c374f3e33280b340 | text | foundation_meta | artifacts/research/gt-foundation-2026-10-21/run_meta.json | /revision | 2026-10-21T16:09:31+00:00 (file time) | 08755c5f2897c441bd0bd4c30577e8151fc571e64000bd6625458af4f33c3b1e | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.license | D-054 checkpoint licence | apache-2.0 | text | foundation_meta | artifacts/research/gt-foundation-2026-10-21/run_meta.json | /license | 2026-10-21T16:09:31+00:00 (file time) | 08755c5f2897c441bd0bd4c30577e8151fc571e64000bd6625458af4f33c3b1e | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.origins | D-054 origins scored | 39761 | origins | foundation_meta | artifacts/research/gt-foundation-2026-10-21/run_meta.json | /origins | 2026-10-21T16:09:31+00:00 (file time) | 08755c5f2897c441bd0bd4c30577e8151fc571e64000bd6625458af4f33c3b1e | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
inv.cgm | CGMacros participants imported | 45 | people | inventory | reports/dataset_inventory.json | /datasets[dataset_id=cgmacros]/parsed_participants | 2026-10-12T13:51:13+00:00 | b5f5be6fc38ef245a4dd8bc133d1f87f46c746c36517c02b58fff7d77ae90b88 | data import | public in repo (tracked)
inv.sh | ShanghaiT2DM participants imported | 100 | people | inventory | reports/dataset_inventory.json | /datasets[dataset_id=shanghai_t2dm]/parsed_participants | 2026-10-12T13:51:13+00:00 | b5f5be6fc38ef245a4dd8bc133d1f87f46c746c36517c02b58fff7d77ae90b88 | data import | public in repo (tracked)
inv.cgm_eval | CGMacros participants evaluated | 14 | people | inventory | reports/dataset_inventory.json | /assessed_participants_by_dataset/cgmacros | 2026-10-12T13:51:13+00:00 | b5f5be6fc38ef245a4dd8bc133d1f87f46c746c36517c02b58fff7d77ae90b88 | data import | public in repo (tracked)
inv.sh_eval | ShanghaiT2DM participants evaluated | 98 | people | inventory | reports/dataset_inventory.json | /assessed_participants_by_dataset/shanghai_t2dm | 2026-10-12T13:51:13+00:00 | b5f5be6fc38ef245a4dd8bc133d1f87f46c746c36517c02b58fff7d77ae90b88 | data import | public in repo (tracked)
reg.serving | Registry serving model, recorded data | lightgbm.dyn-prof-v2.gt-primary-2026-10-13-fusion | text | registry | artifacts/registry.json | /serving/dataset_replay | 2026-10-13T09:21:28+00:00 (file time) | 04540adc4f16226c30886eecd8a6bb199bb89ab7ca08a8f88dfa3d4694b4af1c | registry | local / reproducible-only (git-ignored)
reg.synthetic | Registry serving model, synthetic mode | synthetic-demo.lightgbm.dyn-v2.v1 | text | registry | artifacts/registry.json | /serving/synthetic_demo | 2026-10-13T09:21:28+00:00 (file time) | 04540adc4f16226c30886eecd8a6bb199bb89ab7ca08a8f88dfa3d4694b4af1c | registry | local / reproducible-only (git-ignored)
reg.synthetic_fitted | Registry: synthetic demo model fitted on fictional data | TRUE | flag | registry | artifacts/registry.json | /models[model_version=synthetic-demo.lightgbm.dyn-v2.v1]/synthetic_fitted | 2026-10-13T09:21:28+00:00 (file time) | 04540adc4f16226c30886eecd8a6bb199bb89ab7ca08a8f88dfa3d4694b4af1c | registry | local / reproducible-only (git-ignored)
ver.pytest.passed | Final ledger (f073a10): backend pytest (complete), passed | 221 | tests | ledger | reports/verification/final_2026-10-22_145326Z/ledger.json | /entries[0]/counts/passed | 2026-10-22T14:57:27+00:00 | d2ff869f33cf502aec76d7f20261b8030457aa9b82efae03b478c139169e1447 | scripts/verification_ledger.py | public in repo (tracked)
ver.pytest.skipped | Final ledger (f073a10): backend pytest (complete), skipped | 2 | tests | ledger | reports/verification/final_2026-10-22_145326Z/ledger.json | /entries[0]/counts/skipped | 2026-10-22T14:57:27+00:00 | d2ff869f33cf502aec76d7f20261b8030457aa9b82efae03b478c139169e1447 | scripts/verification_ledger.py | public in repo (tracked)
ver.pytest.failed | Final ledger (f073a10): backend pytest (complete), failed | 0 | tests | ledger | reports/verification/final_2026-10-22_145326Z/ledger.json | /entries[0]/counts/failed | 2026-10-22T14:57:27+00:00 | d2ff869f33cf502aec76d7f20261b8030457aa9b82efae03b478c139169e1447 | scripts/verification_ledger.py | public in repo (tracked)
ver.vitest.passed | Final ledger (f073a10): frontend unit (vitest), passed | 141 | tests | ledger | reports/verification/final_2026-10-22_145326Z/ledger.json | /entries[3]/counts/passed | 2026-10-22T14:57:27+00:00 | d2ff869f33cf502aec76d7f20261b8030457aa9b82efae03b478c139169e1447 | scripts/verification_ledger.py | public in repo (tracked)
ver.vitest.skipped | Final ledger (f073a10): frontend unit (vitest), skipped | 0 | tests | ledger | reports/verification/final_2026-10-22_145326Z/ledger.json | /entries[3]/counts/skipped | 2026-10-22T14:57:27+00:00 | d2ff869f33cf502aec76d7f20261b8030457aa9b82efae03b478c139169e1447 | scripts/verification_ledger.py | public in repo (tracked)
ver.vitest.failed | Final ledger (f073a10): frontend unit (vitest), failed | 0 | tests | ledger | reports/verification/final_2026-10-22_145326Z/ledger.json | /entries[3]/counts/failed | 2026-10-22T14:57:27+00:00 | d2ff869f33cf502aec76d7f20261b8030457aa9b82efae03b478c139169e1447 | scripts/verification_ledger.py | public in repo (tracked)
ver.e2e.passed | Final ledger (f073a10): e2e normal + recorded judge journeys, passed | 29 | tests | ledger | reports/verification/final_2026-10-22_145326Z/ledger.json | /entries[5]/counts/passed | 2026-10-22T14:57:27+00:00 | d2ff869f33cf502aec76d7f20261b8030457aa9b82efae03b478c139169e1447 | scripts/verification_ledger.py | public in repo (tracked)
ver.e2e.skipped | Final ledger (f073a10): e2e normal + recorded judge journeys, skipped | 0 | tests | ledger | reports/verification/final_2026-10-22_145326Z/ledger.json | /entries[5]/counts/skipped | 2026-10-22T14:57:27+00:00 | d2ff869f33cf502aec76d7f20261b8030457aa9b82efae03b478c139169e1447 | scripts/verification_ledger.py | public in repo (tracked)
ver.e2e.failed | Final ledger (f073a10): e2e normal + recorded judge journeys, failed | 0 | tests | ledger | reports/verification/final_2026-10-22_145326Z/ledger.json | /entries[5]/counts/failed | 2026-10-22T14:57:27+00:00 | d2ff869f33cf502aec76d7f20261b8030457aa9b82efae03b478c139169e1447 | scripts/verification_ledger.py | public in repo (tracked)
ver.e2e_public.passed | Final ledger (f073a10): e2e public mode (local, synthetic), passed | 4 | tests | ledger | reports/verification/final_2026-10-22_145326Z/ledger.json | /entries[6]/counts/passed | 2026-10-22T14:57:27+00:00 | d2ff869f33cf502aec76d7f20261b8030457aa9b82efae03b478c139169e1447 | scripts/verification_ledger.py | public in repo (tracked)
ver.e2e_public.skipped | Final ledger (f073a10): e2e public mode (local, synthetic), skipped | 0 | tests | ledger | reports/verification/final_2026-10-22_145326Z/ledger.json | /entries[6]/counts/skipped | 2026-10-22T14:57:27+00:00 | d2ff869f33cf502aec76d7f20261b8030457aa9b82efae03b478c139169e1447 | scripts/verification_ledger.py | public in repo (tracked)
ver.e2e_public.failed | Final ledger (f073a10): e2e public mode (local, synthetic), failed | 0 | tests | ledger | reports/verification/final_2026-10-22_145326Z/ledger.json | /entries[6]/counts/failed | 2026-10-22T14:57:27+00:00 | d2ff869f33cf502aec76d7f20261b8030457aa9b82efae03b478c139169e1447 | scripts/verification_ledger.py | public in repo (tracked)
ver.e2e_summary.passed | Final ledger (f073a10): e2e summary/export, passed | 3 | tests | ledger | reports/verification/final_2026-10-22_145326Z/ledger.json | /entries[7]/counts/passed | 2026-10-22T14:57:27+00:00 | d2ff869f33cf502aec76d7f20261b8030457aa9b82efae03b478c139169e1447 | scripts/verification_ledger.py | public in repo (tracked)
ver.e2e_summary.skipped | Final ledger (f073a10): e2e summary/export, skipped | 0 | tests | ledger | reports/verification/final_2026-10-22_145326Z/ledger.json | /entries[7]/counts/skipped | 2026-10-22T14:57:27+00:00 | d2ff869f33cf502aec76d7f20261b8030457aa9b82efae03b478c139169e1447 | scripts/verification_ledger.py | public in repo (tracked)
ver.e2e_summary.failed | Final ledger (f073a10): e2e summary/export, failed | 0 | tests | ledger | reports/verification/final_2026-10-22_145326Z/ledger.json | /entries[7]/counts/failed | 2026-10-22T14:57:27+00:00 | d2ff869f33cf502aec76d7f20261b8030457aa9b82efae03b478c139169e1447 | scripts/verification_ledger.py | public in repo (tracked)
ver.commit | Tested commit (final ledger) | f073a10f02508f35145b8182bbae13127f7a479e | text | ledger | reports/verification/final_2026-10-22_145326Z/ledger.json | /snapshots[label=before]/revision/commit | 2026-10-22T14:57:27+00:00 | d2ff869f33cf502aec76d7f20261b8030457aa9b82efae03b478c139169e1447 | scripts/verification_ledger.py | public in repo (tracked)
src.cgm.doi | cgmacros DOI | https://doi.org/10.13026/3z8q-x658 | text | sources_yaml | configs/sources.yaml | sources.cgmacros.doi | 2026-10-12T19:30:15+05:30 | a36a4273ba8f592658d4c4c84cb52e7923aa611dc5d6a5dd5f9c114c33fef390 | source manifest | public in repo (tracked)
src.cgm.license | cgmacros licence | CC BY-NC-SA 4.0 | text | sources_yaml | configs/sources.yaml | sources.cgmacros.license | 2026-10-12T19:30:15+05:30 | a36a4273ba8f592658d4c4c84cb52e7923aa611dc5d6a5dd5f9c114c33fef390 | source manifest | public in repo (tracked)
src.cgm.version | cgmacros release | 1.0.0 | text | sources_yaml | configs/sources.yaml | sources.cgmacros.release | 2026-10-12T19:30:15+05:30 | a36a4273ba8f592658d4c4c84cb52e7923aa611dc5d6a5dd5f9c114c33fef390 | source manifest | public in repo (tracked)
src.sh.doi | shanghai_t2dm DOI | https://doi.org/10.6084/m9.figshare.21600933.v5 | text | sources_yaml | configs/sources.yaml | sources.shanghai_t2dm.doi | 2026-10-12T19:30:15+05:30 | a36a4273ba8f592658d4c4c84cb52e7923aa611dc5d6a5dd5f9c114c33fef390 | source manifest | public in repo (tracked)
src.sh.license | shanghai_t2dm licence | CC BY 4.0 | text | sources_yaml | configs/sources.yaml | sources.shanghai_t2dm.license | 2026-10-12T19:30:15+05:30 | a36a4273ba8f592658d4c4c84cb52e7923aa611dc5d6a5dd5f9c114c33fef390 | source manifest | public in repo (tracked)
src.sh.version | shanghai_t2dm release | figshare 21600933 v5 (2023-10-30) | text | sources_yaml | configs/sources.yaml | sources.shanghai_t2dm.release | 2026-10-12T19:30:15+05:30 | a36a4273ba8f592658d4c4c84cb52e7923aa611dc5d6a5dd5f9c114c33fef390 | source manifest | public in repo (tracked)
src.cgm.publisher_n | CGMacros participants per the publisher | 45 | people | sources_yaml | configs/sources.yaml | sources.cgmacros.publisher_participants | 2026-10-12T19:30:15+05:30 | a36a4273ba8f592658d4c4c84cb52e7923aa611dc5d6a5dd5f9c114c33fef390 | source manifest | public in repo (tracked)
cfg.threshold | Event threshold (strictly greater than) | 180 | mg/dL | experiment_yaml | configs/experiment.yaml | events.threshold_mg_dl | 2026-10-12T20:41:34+05:30 | 3e640522905f126ded7524679c0b8bf53709b7d5f8c6e329624e3994d00703df | experiment config | public in repo (tracked)
cfg.confirm | Minutes above threshold to confirm an episode | 15 | min | experiment_yaml | configs/experiment.yaml | events.confirm_minutes | 2026-10-12T20:41:34+05:30 | 3e640522905f126ded7524679c0b8bf53709b7d5f8c6e329624e3994d00703df | experiment config | public in repo (tracked)
cfg.stale | Stale after this many minutes without a reading | 15 | min | experiment_yaml | configs/experiment.yaml | features.stale_threshold_minutes | 2026-10-12T20:41:34+05:30 | 3e640522905f126ded7524679c0b8bf53709b7d5f8c6e329624e3994d00703df | experiment config | public in repo (tracked)
cfg.lookback | Feature lookback window | 180 | min | experiment_yaml | configs/experiment.yaml | features.lookback_minutes | 2026-10-12T20:41:34+05:30 | 3e640522905f126ded7524679c0b8bf53709b7d5f8c6e329624e3994d00703df | experiment config | public in repo (tracked)
cfg.reset_gap | Correction resets after a gap of this length | 180 | min | experiment_yaml | configs/experiment.yaml | adaptation.reset_gap_minutes | 2026-10-12T20:41:34+05:30 | 3e640522905f126ded7524679c0b8bf53709b7d5f8c6e329624e3994d00703df | experiment config | public in repo (tracked)
const.min_per_day | Minutes in a day (definition) | 1440 | min | experiment_yaml | configs/experiment.yaml | constant: minutes per day | 2026-10-12T20:41:34+05:30 | 3e640522905f126ded7524679c0b8bf53709b7d5f8c6e329624e3994d00703df | experiment config | public in repo (tracked)
const.cadence | Dexcom G6 Pro native cadence | 5 | min | sources_yaml | configs/sources.yaml | sources.cgmacros.devices[dexcom_g6_pro].native_cadence_minutes | 2026-10-12T19:30:15+05:30 | a36a4273ba8f592658d4c4c84cb52e7923aa611dc5d6a5dd5f9c114c33fef390 | source manifest | public in repo (tracked)
const.target_cov | Declared prediction-range target coverage | 0.8 | share | experiment_yaml | configs/experiment.yaml | intervals level 0.8 (declared target coverage) | 2026-10-12T20:41:34+05:30 | 3e640522905f126ded7524679c0b8bf53709b7d5f8c6e329624e3994d00703df | experiment config | public in repo (tracked)
sel.lightgbm_dyn-prof-v2.selection_mae_30 | Serving experiment selection (validation segment, population model): lightgbm:dyn-prof-v2, validation MAE 30 | 13.4694 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=lightgbm:dyn-prof-v2]/selection_mae_30 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.lightgbm_dyn-prof-v2.selection_mae_60 | Serving experiment selection (validation segment, population model): lightgbm:dyn-prof-v2, validation MAE 60 | 21.2563 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=lightgbm:dyn-prof-v2]/selection_mae_60 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.lightgbm_dyn-prof-v2.selection_score | Serving experiment selection (validation segment, population model): lightgbm:dyn-prof-v2, selection score (mean of 30/60) | 17.3628 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=lightgbm:dyn-prof-v2]/selection_score | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.catboost_dyn-prof-v2.selection_mae_30 | Serving experiment selection (validation segment, population model): catboost:dyn-prof-v2, validation MAE 30 | 13.5647 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=catboost:dyn-prof-v2]/selection_mae_30 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.catboost_dyn-prof-v2.selection_mae_60 | Serving experiment selection (validation segment, population model): catboost:dyn-prof-v2, validation MAE 60 | 21.3079 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=catboost:dyn-prof-v2]/selection_mae_60 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.catboost_dyn-prof-v2.selection_score | Serving experiment selection (validation segment, population model): catboost:dyn-prof-v2, selection score (mean of 30/60) | 17.4363 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=catboost:dyn-prof-v2]/selection_score | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.lightgbm_dyn-v2.selection_mae_30 | Serving experiment selection (validation segment, population model): lightgbm:dyn-v2, validation MAE 30 | 13.7642 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=lightgbm:dyn-v2]/selection_mae_30 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.lightgbm_dyn-v2.selection_mae_60 | Serving experiment selection (validation segment, population model): lightgbm:dyn-v2, validation MAE 60 | 21.7269 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=lightgbm:dyn-v2]/selection_mae_60 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.lightgbm_dyn-v2.selection_score | Serving experiment selection (validation segment, population model): lightgbm:dyn-v2, selection score (mean of 30/60) | 17.7456 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=lightgbm:dyn-v2]/selection_score | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.catboost_dyn-v2.selection_mae_30 | Serving experiment selection (validation segment, population model): catboost:dyn-v2, validation MAE 30 | 13.8834 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=catboost:dyn-v2]/selection_mae_30 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.catboost_dyn-v2.selection_mae_60 | Serving experiment selection (validation segment, population model): catboost:dyn-v2, validation MAE 60 | 21.6646 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=catboost:dyn-v2]/selection_mae_60 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.catboost_dyn-v2.selection_score | Serving experiment selection (validation segment, population model): catboost:dyn-v2, selection score (mean of 30/60) | 17.774 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=catboost:dyn-v2]/selection_score | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.ridge_dyn-v2.selection_mae_30 | Serving experiment selection (validation segment, population model): ridge:dyn-v2, validation MAE 30 | 14.1273 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=ridge:dyn-v2]/selection_mae_30 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.ridge_dyn-v2.selection_mae_60 | Serving experiment selection (validation segment, population model): ridge:dyn-v2, validation MAE 60 | 22.44 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=ridge:dyn-v2]/selection_mae_60 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.ridge_dyn-v2.selection_score | Serving experiment selection (validation segment, population model): ridge:dyn-v2, selection score (mean of 30/60) | 18.2837 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=ridge:dyn-v2]/selection_score | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.ridge_dyn-prof-v2.selection_mae_30 | Serving experiment selection (validation segment, population model): ridge:dyn-prof-v2, validation MAE 30 | 14.2132 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=ridge:dyn-prof-v2]/selection_mae_30 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.ridge_dyn-prof-v2.selection_mae_60 | Serving experiment selection (validation segment, population model): ridge:dyn-prof-v2, validation MAE 60 | 22.6073 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=ridge:dyn-prof-v2]/selection_mae_60 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.ridge_dyn-prof-v2.selection_score | Serving experiment selection (validation segment, population model): ridge:dyn-prof-v2, selection score (mean of 30/60) | 18.4103 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=ridge:dyn-prof-v2]/selection_score | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.linear_trend_cgm-only-v1.selection_mae_30 | Serving experiment selection (validation segment, population model): linear_trend:cgm-only-v1, validation MAE 30 | 15.6578 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=linear_trend:cgm-only-v1]/selection_mae_30 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.linear_trend_cgm-only-v1.selection_mae_60 | Serving experiment selection (validation segment, population model): linear_trend:cgm-only-v1, validation MAE 60 | 26.3316 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=linear_trend:cgm-only-v1]/selection_mae_60 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.linear_trend_cgm-only-v1.selection_score | Serving experiment selection (validation segment, population model): linear_trend:cgm-only-v1, selection score (mean of 30/60) | 20.9947 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=linear_trend:cgm-only-v1]/selection_score | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.persistence_cgm-only-v1.selection_mae_30 | Serving experiment selection (validation segment, population model): persistence:cgm-only-v1, validation MAE 30 | 16.5101 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=persistence:cgm-only-v1]/selection_mae_30 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.persistence_cgm-only-v1.selection_mae_60 | Serving experiment selection (validation segment, population model): persistence:cgm-only-v1, validation MAE 60 | 26.3232 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=persistence:cgm-only-v1]/selection_mae_60 | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.persistence_cgm-only-v1.selection_score | Serving experiment selection (validation segment, population model): persistence:cgm-only-v1, selection score (mean of 30/60) | 21.4166 | mg/dL | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/candidates[label=persistence:cgm-only-v1]/selection_score | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.selected | Selected candidate (fusion-constrained rule) | lightgbm:dyn-prof-v2 | text | fusion | artifacts/experiments/gt-primary-2026-10-13-fusion/metrics.json | /selection/selected_label | 2026-10-13T08:02:25+00:00 | ec8d2e7aedb517e73fdb75284a8dea24bdb8de4b4f98257c6186d2cea56152d9 | protocol configs/experiments/gt-primary-2026-10-13-fusion.yaml (first commit b60d149 2026-10-13T13:32:50+05:30) | local / reproducible-only (git-ignored)
sel.v2_selected | Selected candidate without the fusion constraint (v2) | lightgbm:dyn-v2 | text | fusion_v2 | artifacts/experiments/gt-primary-2026-10-11-v2/metrics.json | /selection/selected_label | 2026-10-11T16:02:37+00:00 | 895f13b8668618a2c26c81515bc1da30fc6415392ee154f4fa14b207280bbfc8 | protocol configs/experiments/gt-primary-2026-10-11-v1.yaml (first commit aee2b41 2026-10-12T20:37:21+05:30) | local / reproducible-only (git-ignored)
fd.fp30.mean | D-054 foundation minus persistence, 30 min: mean | -0.79313 | mg/dL | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/30/FM_vs_persistence/diff_macro/mean | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.fp30.lo | D-054 foundation minus persistence, 30 min: lo | -1.0092 | mg/dL | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/30/FM_vs_persistence/diff_macro/lo | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.fp30.hi | D-054 foundation minus persistence, 30 min: hi | -0.59145 | mg/dL | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/30/FM_vs_persistence/diff_macro/hi | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.fp30.people_better | D-054 foundation minus persistence, 30 min: people_better | 14 | people | foundation | artifacts/research/gt-foundation-2026-10-21/summary.json | /horizons/30/FM_vs_persistence/diff_macro/people_better | 2026-10-21T16:09:47+00:00 | ebf2a2e6cb4bc2dcba531d23fdad6fad71f4e849933ee9537690bb6f939cf4e7 | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)
fd.params | D-054 checkpoint parameter count | 8652672 | parameters | foundation_meta | artifacts/research/gt-foundation-2026-10-21/run_meta.json | /parameters | 2026-10-21T16:09:31+00:00 (file time) | 08755c5f2897c441bd0bd4c30577e8151fc571e64000bd6625458af4f33c3b1e | protocol configs/experiments/gt-foundation-2026-10-21.yaml (first commit 1967c25 2026-10-21T21:31:06+05:30) | local / reproducible-only (git-ignored)

## Cohorts
Cohorts — who each result counts |  |  |  |  | ← About / navigation
Separate evaluation groups; never compare rows across groups.
Evaluation group | Dataset / version | Participants | What the participant count means | Cadence | Split roles | Horizon (min) | Origins / forecasts | Hours | Onsets | Inspected / unseen status | Opening disclosure | Model scope | Source IDs
Serving test (final test segment) | CGMacros 1.0.0 |  | Type 2 group (HbA1c > 6.4 %), same people as training | 5 min | Chronological 60/10/10/20 within each person: training / selection / calibration / test, purge gaps | 30, 60 |  |  |  | Inspected: the v1/v2 test segment was read before the fusion constraint was declared (retrospective) | Later recordings of the same 14 people — never 'unseen patients' | Serving fusion LightGBM + correction | fus.participants; fus.seg.test.*
Serving model selection (validation segment) | CGMacros 1.0.0 |  | Same 14 people | 5 min | Validation segment used for model, threshold and adaptation choice | 30, 60 |  |  |  | Used for selection | — | Serving candidates | fus.seg.validation.*
Serving calibration segment | CGMacros 1.0.0 |  | Same 14 people | 5 min | Split-conformal range calibration | 30, 60 |  |  |  | Used for calibration | — | Serving ranges | fus.seg.calibration.*
Excluded CGMacros participants | CGMacros 1.0.0 |  | Pre-diabetes or no diabetes; replay only | 5 min | Never used in training or evaluation | — | — | — | — | Not evaluated | — | Replay only | fus.excluded; inv.cgm
Participant-held-out internal (D-032 nested; D-043…D-053 reuse the folds) | CGMacros 1.0.0 |  | Each of the 14 people held out entirely once (5 outer folds) | 5 min | Outer folds assess; inner folds select; calibration people for ranges | 30, 60 (120 in D-043) | — | — | — | Previously inspected people: internal validation, not an untouched cohort | Each study opens its outer folds once behind a lock file | Research procedures (not served) | inv.cgm_eval; nest.*
External (D-032 first opening; D-045 second opening) | ShanghaiT2DM, figshare 21600933 v5 |  | Type 2 people imported (100) with usable recordings | 15 min | Assessment only; nothing chosen on this cohort | 30, 60, 120 | — | — | — | Reserved external cohort | Opened twice (D-032 30/60; D-045 120) | CGM-only 15-min LightGBM — not the fusion model | inv.sh; inv.sh_eval
Foundation baseline (D-054) | CGMacros 1.0.0 |  | All 14 Type 2 people, all eligible origins | 5 min | No training (zero-shot) | 30, 60 |  | — | — | Previously inspected people | — | Chronos-Bolt-Tiny, CGM only | fd.origins
Synthetic demonstration | Fictional synthetic participants | — | Fictional; labelled synthetic everywhere | 5 min | — | 30, 60 | — | — | — | Never evidence of performance | — | synthetic-demo.lightgbm.dyn-v2.v1 | reg.synthetic; reg.synthetic_fitted

## Forecasts
Forecasts — errors by evaluation, model and horizon |  |  |  |  | ← About / navigation
MAE = mean absolute error, macro (equal weight per participant). Difference = candidate − persistence on the same origins (negative = lower error). Saved paired differences and CIs are participant-bootstrap values from the artifacts.
Evaluation group | Model / procedure | Horizon (min) | Aggregation | MAE (mg/dL) | Persistence MAE, same origins | Candidate − persistence (formula) | Relative change % (formula) | RMSE (mg/dL) | Saved paired diff | CI 95 % low | CI 95 % high | Note | Source IDs
Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | Fusion LightGBM + correction (SERVING) | 30 | macro per participant | [8 empty] | Retrospective; same 14 people as training | fus.serv.mae30; fus.pers.mae30; fus.serv.rmse30; fus.cmp.sp.d30; fus.cmp.sp.lo30; fus.cmp.sp.hi30
Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | Fusion LightGBM + correction (SERVING) | 60 | macro per participant | [8 empty] | Retrospective; same 14 people as training | fus.serv.mae60; fus.pers.mae60; fus.serv.rmse60; fus.cmp.sp.d60; fus.cmp.sp.lo60; fus.cmp.sp.hi60
Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | CatBoost + profile + correction (not selected) | 30 | macro per participant | [5 empty] | — | — | — |  | fus.cbpa.mae30; fus.pers.mae30; fus.cbpa.rmse30
Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | CatBoost + profile + correction (not selected) | 60 | macro per participant | [5 empty] | — | — | — |  | fus.cbpa.mae60; fus.pers.mae60; fus.cbpa.rmse60
Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | LightGBM + profile, population | 30 | macro per participant | [5 empty] | — | — | — |  | fus.lgbp.mae30; fus.pers.mae30; fus.lgbp.rmse30
Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | LightGBM + profile, population | 60 | macro per participant | [5 empty] | — | — | — |  | fus.lgbp.mae60; fus.pers.mae60; fus.lgbp.rmse60
Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | LightGBM dynamic-only + correction (v2) | 30 | macro per participant | [5 empty] | — | — | — |  | fus.v2a.mae30; fus.pers.mae30; fus.v2a.rmse30
Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | LightGBM dynamic-only + correction (v2) | 60 | macro per participant | [5 empty] | — | — | — |  | fus.v2a.mae60; fus.pers.mae60; fus.v2a.rmse60
Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | Ridge dynamic + profile | 30 | macro per participant | [5 empty] | — | — | — |  | fus.ridgep.mae30; fus.pers.mae30; fus.ridgep.rmse30
Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | Ridge dynamic + profile | 60 | macro per participant | [5 empty] | — | — | — |  | fus.ridgep.mae60; fus.pers.mae60; fus.ridgep.rmse60
Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | Linear trend | 30 | macro per participant | [5 empty] | — | — | — |  | fus.trend.mae30; fus.pers.mae30; fus.trend.rmse30
Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | Linear trend | 60 | macro per participant | [5 empty] | — | — | — |  | fus.trend.mae60; fus.pers.mae60; fus.trend.rmse60
Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | Persistence | 30 | macro per participant | [5 empty] | — | — | — |  | fus.pers.mae30; fus.pers.mae30; fus.pers.rmse30
Serving model, final test segment (later recordings of the same 14 people; retrospective, inspected) | Persistence | 60 | macro per participant | [5 empty] | — | — | — |  | fus.pers.mae60; fus.pers.mae60; fus.pers.rmse60
Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) (D-032) | Selected research procedure (Ridge CGM-only in 4/5 folds) | 30 | macro per participant |  |  |  |  | not in source | — | — | — | Not served | nest.sel.mae30; nest.pers.mae30
Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) (D-032) | Selected research procedure (Ridge CGM-only in 4/5 folds) | 60 | macro per participant |  |  |  |  | not in source | — | — | — | Not served | nest.sel.mae60; nest.pers.mae60
Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) (D-032) | v2 recipe refit per fold | 30 | macro per participant |  |  |  |  | not in source | — | — | — | Not served | nest.v2.mae30; nest.pers.mae30
Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) (D-032) | v2 recipe refit per fold | 60 | macro per participant |  |  |  |  | not in source | — | — | — | Not served | nest.v2.mae60; nest.pers.mae60
External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) (D-032) | 15-min LightGBM, CGM only | 30 | macro per participant |  |  |  |  | not in source |  |  |  | Not the fusion model | ext.cand.mae30; ext.pers.mae30; ext.cmp.mean_diff30; ext.cmp.lo30; ext.cmp.hi30
External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) (D-032) | 15-min LightGBM, CGM only | 60 | macro per participant |  |  |  |  | not in source |  |  |  | Not the fusion model | ext.cand.mae60; ext.pers.mae60; ext.cmp.mean_diff60; ext.cmp.lo60; ext.cmp.hi60
Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) (D-043) | Two-hour selected procedure, personalised | 30 | macro per participant |  |  |  |  | not in source | — | — | — | Not served | 2h.sel.mae30; 2h.pers.mae30
External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) (D-045) | 15-min LightGBM refit for 120 min | 30 | macro per participant |  |  |  |  | not in source | — | — | — | Second opening of the cohort | 2hx.mae30; 2hx.pers30
Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) (D-043) | Two-hour selected procedure, personalised | 60 | macro per participant |  |  |  |  | not in source | — | — | — | Not served | 2h.sel.mae60; 2h.pers.mae60
External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) (D-045) | 15-min LightGBM refit for 120 min | 60 | macro per participant |  |  |  |  | not in source | — | — | — | Second opening of the cohort | 2hx.mae60; 2hx.pers60
Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) (D-043) | Two-hour selected procedure, personalised | 120 | macro per participant |  |  |  |  | not in source |  |  |  | Not served | 2h.sel.mae120; 2h.pers.mae120; 2h.cmp.d; 2h.cmp.lo; 2h.cmp.hi
External ShanghaiT2DM, 98 people, CGM-only procedure (not the serving fusion model) (D-045) | 15-min LightGBM refit for 120 min | 120 | macro per participant |  |  |  |  | not in source |  |  |  | Second opening of the cohort | 2hx.mae120; 2hx.pers120; 2hx.cmp.mean_diff; 2hx.cmp.lo; 2hx.cmp.hi
Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) (D-053) | Constrained selection S (not adopted) | 30 | macro per participant |  |  |  |  | not in source | — | — | — | Gate failed; not adopted | cs.S.mae30; cs.persistence.mae30
Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) (D-053) | Serving recipe refitted in the same folds (R0) | 30 | macro per participant |  |  |  |  | not in source | — | — | — |  | cs.R0.mae30; cs.persistence.mae30
Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) — all origins (D-054) | Chronos-Bolt-Tiny zero-shot, CGM only | 30 | macro per participant |  |  |  |  | not in source | — | — | — | Not fine-tuned, not served | fd.FM30; fd.P30
Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) (D-053) | Constrained selection S (not adopted) | 60 | macro per participant |  |  |  |  | not in source | — | — | — | Gate failed; not adopted | cs.S.mae60; cs.persistence.mae60
Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) (D-053) | Serving recipe refitted in the same folds (R0) | 60 | macro per participant |  |  |  |  | not in source |  |  |  | Saved diff on this row = S − R0 | cs.R0.mae60; cs.persistence.mae60; cs.p60.mean_diff; cs.p60.lo; cs.p60.hi
Participant-held-out internal (14 CGMacros people, each held out once; previously inspected) — all origins (D-054) | Chronos-Bolt-Tiny zero-shot, CGM only | 60 | macro per participant |  |  |  |  | not in source |  |  |  | Not fine-tuned, not served | fd.FM60; fd.P60; fd.fp.mean; fd.fp.lo; fd.fp.hi
Chart data: serving test MAE vs persistence (references the table above)
Horizon | Serving fusion LightGBM | Persistence
30 min
60 min

## WarningRanges
Warnings and prediction ranges |  |  |  |  | ← About / navigation
Numerator and denominator are shown for every share. Notice shares use ALL onsets; the median lead uses DETECTED onsets. False-alarm rates divide by the saved observed / eligible patient-day denominators (not by rounded hours).
Section | Metric | Value (formula) | Unit | Numerator | Denominator | Definition / denominator meaning | Saved value to agree with | Agrees? | Source IDs
Serving warning (test) | Onsets (eligible new sustained > 180 episodes) |  | onsets | — | — | Distinct new onsets; not origins | — | — | fus.ev.true_onsets
Serving warning (test) | Detected onsets |  | onsets | — | — | Onsets matched to a warning episode | — | — | fus.ev.detected_onsets
Serving warning (test) | Missed onsets |  | onsets | — | — |  | — | — | fus.ev.true_onsets; fus.ev.detected_onsets
Serving warning (test) | Detection share (detected / all onsets) |  | share |  |  | Recall over all onsets | — | — | fus.ev.detected_onsets; fus.ev.true_onsets
Serving warning (test) | Warned ≥ 15 min ahead, share of ALL onsets |  | share |  |  | All onsets; a missed onset counts against it | — | — | wu.ge15; wu.onsets
Serving warning (test) | Warned ≥ 30 min ahead, share of ALL onsets |  | share |  |  | All onsets (headline 48.4 %) | — | — | wu.ge30; wu.onsets
Serving warning (test) | Warned ≥ 60 min ahead, share of ALL onsets |  | share |  |  | All onsets | — | — | wu.ge60; wu.onsets
Serving warning (test) | Median qualified lead (DETECTED onsets only) |  | min | — |  | Denominator = detected onsets with a qualified lead; not comparable to the all-onset shares |  |  | fus.ev.qualified_lead_median; fus.ev.qualified_lead_n; wu.lead_median
Serving warning (test) | Qualified lead IQR low (detected onsets) |  | min | — | — |  |  |  | fus.ev.iqr_lo; wu.lead_q25
Serving warning (test) | Qualified lead IQR high (detected onsets) |  | min | — | — |  |  |  | fus.ev.iqr_hi; wu.lead_q75
Serving warning (test) | False-alarm episodes |  | episodes | — | — |  | — | — | fus.ev.false_alarm_episodes
Serving warning (test) | Warning episodes |  | episodes | — | — |  | — | — | fus.ev.warning_episodes
Serving warning (test) | Precision (true episodes / episodes) |  | ratio | — |  |  |  |  | fus.ev.warning_episodes; fus.ev.false_alarm_episodes; fus.ev.precision
Serving warning (test) | False alarms per OBSERVED patient-day |  | per day |  |  | Observed patient-days from saved event metadata (time base: issued origins × 5 min) |  |  | fus.ev.false_alarm_episodes; fus.ev.observed_patient_days; fus.ev.false_alarms_per_observed_patient_day
Serving warning (test) | False alarms per ELIGIBLE patient-day |  | per day |  |  | Eligible patient-days (below-threshold origins) — a different, smaller base |  |  | fus.ev.false_alarm_episodes; fus.ev.eligible_patient_days; fus.ev.false_alarms_per_patient_day
Serving warning (test) | Time in warning (share of eligible origins) |  | share |  |  | Post-hoc diagnostic |  |  | wu.firing; wu.eligible; wu.share_eligible
Serving warning (test) | Warning margin (score = max forecast − 180) |  | mg/dL | — | — | Chosen on validation by event F1 | — | — | fus.ev.threshold
Persistence warning (test) | False alarms per observed patient-day |  | per day | — | — | Fires at 150 mg/dL | — | — | fus.pev.false_alarms_per_observed_patient_day
Serving ranges (test) | 80 % range coverage, 30 min |  | share | — |  | Pooled over forecasts | — | — | fus.iv.coverage30; fus.iv.n30
Serving ranges (test) | 80 % range coverage, 60 min |  | share | — |  | Pooled over forecasts | — | — | fus.iv.coverage60; fus.iv.n60
Serving ranges (test) | Coverage gap vs target, 30 min (points) |  | points | — | — | Negative = under-coverage | — | — | fus.iv.coverage30; const.target_cov
Serving ranges (test) | Coverage gap vs target, 60 min (points) |  | points | — | — | Negative = under-coverage | — | — | fus.iv.coverage60; const.target_cov
Serving ranges (test) | Per-person coverage min, 30 min |  | share | — | — | Weakest participant | — | — | fus.iv.per_patient_min30
Serving ranges (test) | Per-person coverage max, 30 min |  | share | — | — |  | — | — | fus.iv.per_patient_max30
Serving ranges (test) | Per-person coverage min, 60 min |  | share | — | — | Weakest participant | — | — | fus.iv.per_patient_min60
Serving ranges (test) | Per-person coverage max, 60 min |  | share | — | — |  | — | — | fus.iv.per_patient_max60
Serving ranges (test) | Range width (full), 30 min |  | mg/dL | — | — | Full width: 2 × half-width q | — | — | fus.iv.width30
Serving ranges (test) | Range width (full), 60 min |  | mg/dL | — | — | Full width: 2 × half-width q | — | — | fus.iv.width60
Trend context — VALIDATION | falling_fast, 30 min: macro MAE |  | mg/dL | — |  | Validation segment (selection data), not the test | — | — | tr.falling_fast.mae30; tr.falling_fast.n30
Trend context — VALIDATION | falling_fast, 30 min: served range coverage |  | share | — |  | Validation segment | — | — | tr.falling_fast.cov30; tr.falling_fast.n30
Trend context — VALIDATION | falling_fast, 60 min: macro MAE |  | mg/dL | — |  | Validation segment (selection data), not the test | — | — | tr.falling_fast.mae60; tr.falling_fast.n60
Trend context — VALIDATION | falling_fast, 60 min: served range coverage |  | share | — |  | Validation segment | — | — | tr.falling_fast.cov60; tr.falling_fast.n60
Trend context — VALIDATION | falling, 30 min: macro MAE |  | mg/dL | — |  | Validation segment (selection data), not the test | — | — | tr.falling.mae30; tr.falling.n30
Trend context — VALIDATION | falling, 30 min: served range coverage |  | share | — |  | Validation segment | — | — | tr.falling.cov30; tr.falling.n30
Trend context — VALIDATION | falling, 60 min: macro MAE |  | mg/dL | — |  | Validation segment (selection data), not the test | — | — | tr.falling.mae60; tr.falling.n60
Trend context — VALIDATION | falling, 60 min: served range coverage |  | share | — |  | Validation segment | — | — | tr.falling.cov60; tr.falling.n60
Trend context — VALIDATION | steady, 30 min: macro MAE |  | mg/dL | — |  | Validation segment (selection data), not the test | — | — | tr.steady.mae30; tr.steady.n30
Trend context — VALIDATION | steady, 30 min: served range coverage |  | share | — |  | Validation segment | — | — | tr.steady.cov30; tr.steady.n30
Trend context — VALIDATION | steady, 60 min: macro MAE |  | mg/dL | — |  | Validation segment (selection data), not the test | — | — | tr.steady.mae60; tr.steady.n60
Trend context — VALIDATION | steady, 60 min: served range coverage |  | share | — |  | Validation segment | — | — | tr.steady.cov60; tr.steady.n60
Trend context — VALIDATION | rising, 30 min: macro MAE |  | mg/dL | — |  | Validation segment (selection data), not the test | — | — | tr.rising.mae30; tr.rising.n30
Trend context — VALIDATION | rising, 30 min: served range coverage |  | share | — |  | Validation segment | — | — | tr.rising.cov30; tr.rising.n30
Trend context — VALIDATION | rising, 60 min: macro MAE |  | mg/dL | — |  | Validation segment (selection data), not the test | — | — | tr.rising.mae60; tr.rising.n60
Trend context — VALIDATION | rising, 60 min: served range coverage |  | share | — |  | Validation segment | — | — | tr.rising.cov60; tr.rising.n60
Trend context — VALIDATION | rising_fast, 30 min: macro MAE |  | mg/dL | — |  | Validation segment (selection data), not the test | — | — | tr.rising_fast.mae30; tr.rising_fast.n30
Trend context — VALIDATION | rising_fast, 30 min: served range coverage |  | share | — |  | Validation segment | — | — | tr.rising_fast.cov30; tr.rising_fast.n30
Trend context — VALIDATION | rising_fast, 60 min: macro MAE |  | mg/dL | — |  | Validation segment (selection data), not the test | — | — | tr.rising_fast.mae60; tr.rising_fast.n60
Trend context — VALIDATION | rising_fast, 60 min: served range coverage |  | share | — |  | Validation segment | — | — | tr.rising_fast.cov60; tr.rising_fast.n60
Trend context — FINAL TEST (D-050, read once) | rising, 30 min: served range coverage |  | share | — |  | Final test, retrospective; rising/falling = ¦slope¦ > 1 mg/dL/min | — | — | rg.inc.rising30; rg.inc.risingn30
Trend context — FINAL TEST (D-050, read once) | rising, 60 min: served range coverage |  | share | — |  | Final test, retrospective; rising/falling = ¦slope¦ > 1 mg/dL/min | — | — | rg.inc.rising60; rg.inc.risingn60
Trend context — FINAL TEST (D-050, read once) | steady, 30 min: served range coverage |  | share | — |  | Final test, retrospective; rising/falling = ¦slope¦ > 1 mg/dL/min | — | — | rg.inc.steady30; rg.inc.steadyn30
Trend context — FINAL TEST (D-050, read once) | steady, 60 min: served range coverage |  | share | — |  | Final test, retrospective; rising/falling = ¦slope¦ > 1 mg/dL/min | — | — | rg.inc.steady60; rg.inc.steadyn60
Trend context — FINAL TEST (D-050, read once) | falling, 30 min: served range coverage |  | share | — |  | Final test, retrospective; rising/falling = ¦slope¦ > 1 mg/dL/min | — | — | rg.inc.falling30; rg.inc.fallingn30
Trend context — FINAL TEST (D-050, read once) | falling, 60 min: served range coverage |  | share | — |  | Final test, retrospective; rising/falling = ¦slope¦ > 1 mg/dL/min | — | — | rg.inc.falling60; rg.inc.fallingn60
Chart data: 80 % range coverage by evaluation (target line = 80 %)
Evaluation | Coverage | Target
Serving 30 min (test)
Serving 60 min (test)
Nested research 30 min
Nested research 60 min
D-050 candidate 30 (not served)
D-050 candidate 60 (not served)
D-053 S 30 (not adopted)
D-053 S 60 (not adopted)
Two-hour 120 (D-043)
Two-hour pooled 120 (D-046)
External 120 (D-045)

## Studies
Studies — D-043 onward, including every negative or unopened result |  |  |  |  | ← About / navigation
Protocol-before-run = protocol file's first git commit precedes the saved run/open time (git dates are local commit evidence, not a clinical registration). 'not opened' / 'unavailable' are never zero.
Decision | Experiment ID | Kind | Hypothesis / purpose | Population and role | Protocol file | Protocol first commit (UTC) | Run start / assessment opened (UTC) | Protocol-before-run (formula) | Fit budget | Assessment-opening status | Candidate metric | Control metric | Candidate − control (formula) | Gate outcome (Gates sheet) | Adoption decision | Negative findings / limits | Report | Artifact status
D-032 | gt-nested-2026-10-12-v1 | Nested internal + external | New-participant and external evidence for the procedure | 14 CGMacros people held out once; 98 ShanghaiT2DM external | configs/experiments/gt-nested-2026-10-12-v1.yaml | 2026-10-12 15:10 |  |  | — | Opened once (internal and external) |  |  |  | — | Research candidate, not promoted | Git commit of the protocol file postdates its freeze/open timestamps; the frozen protocol hash in report.json is the evidence | reports/nested_benchmark_report.md | local / reproducible-only (git-ignored)
D-043 | gt-2h-2026-10-20 | Two-hour forecast + warning | Does a 120-min forecast and warning work on people held out? | 14 CGMacros people held out once (inspected) | configs/experiments/gt-2h-2026-10-20.yaml | 2026-10-20 09:20 |  |  | Same procedure as D-032 | Opened once |  |  |  | — | Research only; demo model unchanged | Median lead 15 min; ranges under-cover (76.5 %) | reports/two_hour_benchmark.md | local / reproducible-only (git-ignored)
D-044 | gt-2h-clf-2026-10-20 | Direct onset classifier | Does a classifier trained on the event warn earlier at the same workload? | Folds of D-043 | configs/experiments/gt-2h-clf-2026-10-20.yaml | 2026-10-20 09:57 |  |  | Fixed settings | Opened, then re-opened once (post-hoc diagnostic, reason logged) | — | — |  | — | Rejected (negative result) | Threshold hit the grid minimum; warning active ~99 % of eligible time | reports/two_hour_classifier.md | local / reproducible-only (git-ignored)
D-045 | gt-2h-ext-2026-10-20 | External two-hour | Does the 120-min forecast hold on an independent cohort? | 98 ShanghaiT2DM people (external) | configs/experiments/gt-2h-ext-2026-10-20.yaml | 2026-10-20 10:09 |  |  | 15-min track of D-032 | Second opening of the external cohort |  |  |  | — | Research only | CGM only; not the fusion model | reports/two_hour_external.md | local / reproducible-only (git-ignored)
D-046 | gt-warn-range-2026-10-20 | Re-analysis of saved D-043 predictions | Earlier warning (A), pooled ranges (B), profile prior (C) | Same 14 inspected people | configs/experiments/gt-warn-range-2026-10-20.yaml | 2026-10-20 12:22 |  |  | No model refit | Outer predictions read once |  |  |  | — | B positive (not applied to serving); A and C negative | A: +0.9 pt [−5.5, +6.6]; C: CI crosses zero | reports/warning_range_experiment.md | local / reproducible-only (git-ignored)
D-047 | — | Software (public synthetic mode) | Not a model experiment: opt-in public synthetic demonstration | Synthetic participants only | — | — | — |  | — | not applicable | — | — |  | — | Implemented; not deployed | No performance claim | docs/deploy.md | —
D-048 | gt-blend-2026-10-21 | Saved-prediction blend | Does a LightGBM/CatBoost blend lower 60-min error by ≥ 0.5? | 12 development people (inner out-of-fold) | configs/experiments/gt-blend-2026-10-21.yaml | 2026-10-20 15:07 |  |  | No fit (saved predictions) | NOT OPENED (stop rule: target not met) |  |  |  |  | Target not met; incumbent retained | Gain 0.02 mg/dL vs needed 0.5 | docs/decisions.md D-048 | local / reproducible-only (git-ignored)
D-049 | gt-retrain-2026-10-21 | Retraining arms | Do weighting or history features lower 60-min error by ≥ 0.5? | 12 development people (inner out-of-fold) | configs/experiments/gt-retrain-2026-10-21.yaml | 2026-10-20 17:00 |  |  | 24 fresh fits, fixed settings | NOT OPENED (no arm met the gate) | — | — |  |  | B0 retained | Best arm B1 gain 0.41 < 0.5; descriptive population-only result is after selection | reports/retrain_study.md | local / reproducible-only (git-ignored)
D-050 | gt-ranges-2026-10-21 | Range shape study | Does a trend-stratified range beat the global range before test? | Serving pipeline; validation→calibration (pre-test), test read once | configs/experiments/gt-ranges-2026-10-21.yaml | 2026-10-20 17:40 |  |  | No model refit | Final test read once, retrospective (inspected segment) |  |  |  |  | G kept; pooled calibration is a research candidate only | Candidate coverage is on inspected data — not the served interval | reports/serving_ranges_study.md | local / reproducible-only (git-ignored)
D-051 | gt-tune-2026-10-21 | Nested tuning | Does tuned 24-h history (C1) beat equally tuned current features (C0)? | 14 people held out once (inspected) | configs/experiments/gt-tune-2026-10-21.yaml | 2026-10-20 17:45 |  |  | 30 seeded trials per horizon (implemented budget) | Outer folds opened once |  |  |  |  | Not adopted | Small internal gain vs equally tuned C0 is not a release and is not the same as a comparison with the refitted serving recipe | reports/tune_study.md | local / reproducible-only (git-ignored)
D-052 | gt-joint-2026-10-21 | Joint model + correction selection | Does joint selection on personalised error beat R0? | 14 people held out once (inspected) | configs/experiments/gt-joint-2026-10-21.yaml | 2026-10-20 18:41 |  |  | 1,570 fits | Outer folds opened once |  |  |  |  | Not adopted | High-target +1.70 > 1.0; warning guardrails failed; mixed-fit range erratum | reports/joint_study.md | local / reproducible-only (git-ignored)
D-053 | gt-constrained-2026-10-21 | Constrained (peak-aware) selection | Can selection constrained on high-target error and workload beat R0? | 14 people held out once (inspected) | configs/experiments/gt-constrained-2026-10-21.yaml | 2026-10-21 17:02 |  |  | 1,490 logical / 2,930 physical fits | Assessment opened once |  |  |  |  | NOT ADOPTED (primary gate failed) | High-target worsening +1.0658 vs +1.0 limit | reports/constrained_study.md | local / reproducible-only (git-ignored)
D-054 | gt-foundation-2026-10-21 | Zero-shot foundation baseline | Does a pretrained time-series model beat the tuned tree? | 14 people, all eligible origins (inspected) | configs/experiments/gt-foundation-2026-10-21.yaml | 2026-10-21 16:01 |  |  | No fit (zero-shot); one descriptive residual template | Scored once |  |  |  | — | Not served; descriptive | Not better than refitted recipe; worse above 180; pretraining data may include public CGM | reports/foundation_baseline.md | local / reproducible-only (git-ignored)

## Gates
Gates — the declared rules applied to unrounded saved values |  |  |  |  | ← About / navigation
Value column holds the unrounded source value (display format may round). Result = formula; 'Saved outcome' = the artifact's own flag. A study passes only if every criterion row passes.
Study | Gate group | Criterion (as declared) | Metric value (unrounded) | Comparator | Limit | Result (formula) | Saved outcome | Agrees with saved? | Sign convention / note | Source IDs
D-048 | Development target | Selected alpha improves 60-min dev MAE vs alpha = 1 by ≥ 0.5 mg/dL |  | >= | 0.5 |  | — | — | Improvement positive = gain (alpha = 1 minus selected) | blend.improvement_60
D-048 | Guardrail | 30-min worsening ≤ 0.25 |  | <= | 0.25 |  | — | — | Candidate − control | blend.worsening_30
D-048 | Guardrail | High-glucose 60-min worsening ≤ 1.0 |  | <= | 1 |  | — | — | Candidate − control | blend.high_worsening_60
D-048 | Guardrail | Availability drop ≤ 0.5 points |  | <= | 0.5 |  | — | — |  | blend.availability_drop_points
D-049 | Gate (A1) | A1 60-min personalised gain vs B0 ≥ 0.5 |  | >= | 0.5 |  | — | — | Gain positive = lower error (B0 − arm) | rt.A1.gain_60
D-049 | Gate (A2) | A2 60-min personalised gain vs B0 ≥ 0.5 |  | >= | 0.5 |  | — | — | Gain positive = lower error | rt.A2.gain_60
D-049 | Gate (B1) | B1 60-min personalised gain vs B0 ≥ 0.5 |  | >= | 0.5 |  | — | — | Gain positive = lower error | rt.B1.gain_60
D-050 | Adoption (T replaces G) | Pre-test mean ¦coverage − 0.80¦: T better than G by ≥ 1.0 point |  | >= | 1 |  | — | — | G − T in points; positive = T closer to 80 % | rg.err.G; rg.err.T
D-050 | Adoption (T replaces G) | T mean width ≤ 1.15 × G width |  | <= | 1.15 |  | — | — | Ratio | rg.width.T; rg.width.G
D-051 | Adoption | C1 − C0 at 60 min ≤ −0.5 mg/dL |  | <= | -0.5 |  | — | — | Candidate − control (negative = gain) | tu.c1c0.mean_diff
D-051 | Adoption | Bootstrap 95 % CI upper < 0 |  | < | 0 |  | — | — | Participant bootstrap | tu.c1c0.hi
D-051 | Adoption | 30-min worsening ≤ 0.25 |  | <= | 0.25 |  | — | — | Candidate − control | tu.adopt.worsening_30
D-051 | Adoption | High-glucose 60-min worsening ≤ 1.0 |  | <= | 1 |  | — | — | Candidate − control | tu.adopt.high_worsening_60
D-052 | Primary | S − R0 at 60 min ≤ −0.5 |  | <= | -0.5 |  | — | — | Candidate − control | jt.gate.mean_diff_60
D-052 | Primary | 95 % CI upper < 0 |  | < | 0 |  | — | — |  | jt.gate.ci95_60_hi
D-052 | Primary | 30-min worsening ≤ 0.25 |  | <= | 0.25 |  | — | — |  | jt.gate.worsening_30
D-052 | Primary | High-target (> 180) 60-min worsening ≤ 1.0 |  | <= | 1 |  | — | — | Fails | jt.gate.high_worsening_60
D-052 | Primary | Availability drop ≤ 0.5 points |  | <= | 0.5 |  | — | — |  | jt.gate.availability_drop_points
D-052 | Release guardrail | False alarms/observed day ≤ R0 + 0.10 |  | <= | 0.1 |  | — | — | S − R0 | jt.rel.fa_per_day_delta
D-052 | Release guardrail | Time in warning ≤ R0 + 2.0 points |  | <= | 2 |  | — | — | S − R0 | jt.rel.time_in_warning_delta_points
D-052 | Release guardrail | Share of all onsets warned ≥ 30 min ≥ R0 − 2.0 points |  | >= | -2 |  | — | — | S − R0 | jt.rel.warned30_delta_points
D-053 | Primary | S − R0 at 60 min ≤ −0.5 (inclusive) |  | <= | -0.5 |  |  |  | Candidate − control | cs.pg.mean_diff_60; cs.pgok.mean_diff_60
D-053 | Primary | Paired bootstrap 95 % CI upper < 0 (strict) |  | < | 0 |  |  |  | 2,000 resamples | cs.pg.ci95_60_hi; cs.pgok.ci95_upper_below_0
D-053 | Primary | 30-min worsening ≤ +0.25 (inclusive) |  | <= | 0.25 |  |  |  |  | cs.pg.worsening_30; cs.pgok.worsening_30
D-053 | Primary | High-target (> 180) 60-min worsening ≤ +1.0 (inclusive) |  | <= | 1 |  |  |  | Unrounded ≈ +1.0658 > +1.0 → FAIL at any display rounding | cs.pg.high_worsening_60; cs.pgok.high_worsening_60
D-053 | Primary | Availability drop ≤ 0.5 points |  | <= | 0.5 |  |  |  |  | cs.pg.availability_drop_points; cs.pgok.availability_drop
D-053 | Operational | False alarms/observed day ≤ R0-P + 0.10 |  | <= | 0.1 |  | — | — | S − R0-P | cs.w.S.fa_day; cs.w.R0P.fa_day
D-053 | Operational | Time in warning ≤ R0-P + 2.0 points |  | <= | 2 |  | — | — | S − R0-P | cs.w.S.tiw; cs.w.R0P.tiw
D-053 | Operational | Share of all onsets warned ≥ 30 min ≥ R0-P − 2.0 points |  | >= | -2 |  | — | — | S − R0-P | cs.w.S.share30; cs.w.R0P.share30
D-048 | OVERALL | All criteria above pass | [6 empty] | Not met → incumbent retained; outer folds not opened
D-049 | OVERALL | Any arm passes its gate → that arm qualifies | [6 empty] | No arm met the gate; B0 retained; outer folds not opened
D-050 | OVERALL | Both adoption criteria → T adopted; else G kept | [6 empty] | Saved choice G (served shape kept)
D-051 | OVERALL | All criteria above pass | [6 empty] | Not adopted (small internal gain vs equally tuned C0 is not a release)
D-052 | OVERALL | All criteria above pass | [6 empty] | Not adopted; serving unchanged
D-053 | OVERALL | All criteria above pass | [6 empty] | Primary gate FAILED (high-target); operational passed; NOT ADOPTED
D-053 | OPERATIONAL | Operational criteria pass
D-053 | PRIMARY | Primary point criteria pass
D-054 | — | No adoption gate: descriptive research baseline (F1 residual reported, not selected) |  |  |  | not applicable

## Verification
Verification — what was run, where, and with what outcome |  |  |  |  | ← About / navigation
The latest full local matrix is dated to commit f073a10 (22 Oct); it is not a pass on any later revision. Rows marked author-reported have no saved log in the repository.
Evidence class | Run | Suite | Command (repo-relative) | Tested commit | Data mode | Start (UTC) | End (UTC) | Exit | Status | Total | Passed | Failed | Skipped | Skip reasons | Log | JUnit
Saved local result (final) | final_2026-10-22_145326Z | backend pytest (complete) | uv run pytest -q --junitxml=reports/verification/final_2026-10-22_145326Z/junit/pytest.xml | f073a10f02 | local | 2026-10-22 14:57 | 2026-10-22 14:59 | 0 | passed | 223 | 221 | 0 | 2 | could not import 'torch': No module named 'torch' (1); compared once, from the fusion side (1) | reports/verification/final_2026-10-22_145326Z/logs/pytest.log | reports/verification/final_2026-10-22_145326Z/junit/pytest.xml
Saved local result (final) | final_2026-10-22_145326Z | python lint (ruff; untracked private audit folder excluded) | uv run ruff check --extend-exclude reports/overnight_codex . | f073a10f02 | local | 2026-10-22 14:59 | 2026-10-22 14:59 | 0 | passed | — | — | — | — |  | reports/verification/final_2026-10-22_145326Z/logs/ruff.log
Saved local result (final) | final_2026-10-22_145326Z | frontend typecheck | npm run typecheck | f073a10f02 | local | 2026-10-22 14:59 | 2026-10-22 14:59 | 0 | passed | — | — | — | — |  | reports/verification/final_2026-10-22_145326Z/logs/typecheck.log
Saved local result (final) | final_2026-10-22_145326Z | frontend unit (vitest) | npm test -- --reporter=default --reporter=junit --outputFile.junit=reports/verification/final_2026-10-22_145326Z/junit/vitest.xml | f073a10f02 | local | 2026-10-22 14:59 | 2026-10-22 14:59 | 0 | passed | 141 | 141 | 0 | 0 |  | reports/verification/final_2026-10-22_145326Z/logs/vitest.log | reports/verification/final_2026-10-22_145326Z/junit/vitest.xml
Saved local result (final) | final_2026-10-22_145326Z | frontend production build | npm run build | f073a10f02 | local | 2026-10-22 14:59 | 2026-10-22 14:59 | 0 | passed | — | — | — | — |  | reports/verification/final_2026-10-22_145326Z/logs/build.log
Saved local result (final) | final_2026-10-22_145326Z | e2e normal + recorded judge journeys | npm run e2e -- --workers=1 --reporter=list,junit --output=<scratch>/final/pw/e2e | f073a10f02 | recorded data | 2026-10-22 14:59 | 2026-10-22 15:01 | 0 | passed | 29 | 29 | 0 | 0 |  | reports/verification/final_2026-10-22_145326Z/logs/e2e.log | reports/verification/final_2026-10-22_145326Z/junit/e2e.xml
Saved local result (final) | final_2026-10-22_145326Z | e2e public mode (local, synthetic) | npx playwright test --config playwright.public.config.ts --workers=1 --reporter=list,junit --output=<scratch>/final/pw/public | f073a10f02 | synthetic | 2026-10-22 15:01 | 2026-10-22 15:01 | 0 | passed | 4 | 4 | 0 | 0 |  | reports/verification/final_2026-10-22_145326Z/logs/e2e-public.log | reports/verification/final_2026-10-22_145326Z/junit/e2e-public.xml
Saved local result (final) | final_2026-10-22_145326Z | e2e summary/export | npx playwright test --config playwright.summary.config.ts --workers=1 --reporter=list,junit --output=<scratch>/final/pw/summary | f073a10f02 | local | 2026-10-22 15:01 | 2026-10-22 15:02 | 0 | passed | 3 | 3 | 0 | 0 |  | reports/verification/final_2026-10-22_145326Z/logs/e2e-summary.log | reports/verification/final_2026-10-22_145326Z/junit/e2e-summary.xml
Saved local result (final) | final_2026-10-22_145326Z | selected-claim consistency (evidence manifest) | uv run python scripts/evidence_manifest.py | f073a10f02 | local | 2026-10-22 15:02 | 2026-10-22 15:02 | 0 | passed | — | — | — | — |  | reports/verification/final_2026-10-22_145326Z/logs/evidence_manifest.log
Saved local result (final) | final_2026-10-22_145326Z | README/submission consistency | uv run python scripts/check_submission.py | f073a10f02 | local | 2026-10-22 15:02 | 2026-10-22 15:02 | 0 | passed | — | — | — | — |  | reports/verification/final_2026-10-22_145326Z/logs/check_submission.log
Saved local result (final) | final_2026-10-22_145326Z | D-053 runner validate (read-only, saved artifacts) | uv run python scripts/constrained_study.py validate | f073a10f02 | local | 2026-10-22 15:02 | 2026-10-22 15:02 | 0 | passed | — | — | — | — |  | reports/verification/final_2026-10-22_145326Z/logs/constrained_validate.log
Saved local result (final) | final_2026-10-22_145326Z | serving parity: registered incumbent, new output path | uv run python scripts/serving_parity.py --output reports/verification/final_2026-10-22_145326Z/parity/serving_parity.json --detail-output <scratch>/final/parity/serving_parity_detail.json | f073a10f02 | recorded data | 2026-10-22 15:02 | 2026-10-22 15:02 | 0 | passed | — | — | — | — |  | reports/verification/final_2026-10-22_145326Z/logs/parity.log
Saved local result (parity) | reports/serving_parity.json | App replay vs saved offline test forecasts | uv run python scripts/serving_parity.py | — | recorded data |  |  | — |  |  |  |  | — | tolerance 0.05 mg/dL (API displays 0.1 mg/dL) | reports/serving_parity.json | —
Saved local result (parity) | final ledger parity copy | App replay vs saved offline test forecasts | uv run python scripts/serving_parity.py | — | recorded data |  |  | — |  |  |  |  | — | tolerance 0.05 mg/dL (API displays 0.1 mg/dL) | reports/verification/final_2026-10-22_145326Z/parity/serving_parity.json | —
History (earlier ledger; failures preserved) | overnight | python lint | uv run ruff check . | fd87ca2746 | local | 2026-10-21 16:03 | 2026-10-21 16:03 | 0 | passed | — | — | — | — |  | reports/verification/overnight/baseline/logs/ruff.log
History (earlier ledger; failures preserved) | overnight | backend pytest (complete) | uv run pytest -q --junitxml=.claude/worktrees/qa/reports/verification/overnight/baseline/junit/pytest.xml | fd87ca2746 | local | 2026-10-21 16:03 | 2026-10-21 16:05 | 0 | passed | 148 | 147 | 0 | 1 | compared once, from the fusion side (1) | reports/verification/overnight/baseline/logs/pytest.log | reports/verification/overnight/baseline/junit/pytest.xml
History (earlier ledger; failures preserved) | overnight | frontend typecheck | npm run typecheck | fd87ca2746 | local | 2026-10-21 16:05 | 2026-10-21 16:05 | 0 | passed | — | — | — | — |  | reports/verification/overnight/baseline/logs/typecheck.log
History (earlier ledger; failures preserved) | overnight | frontend unit (vitest) | npm test -- --reporter=default --reporter=junit --outputFile.junit=.claude/worktrees/qa/reports/verification/overnight/baseline/junit/vitest.xml | fd87ca2746 | local | 2026-10-21 16:05 | 2026-10-21 16:05 | 0 | passed | 131 | 131 | 0 | 0 |  | reports/verification/overnight/baseline/logs/vitest.log | reports/verification/overnight/baseline/junit/vitest.xml
History (earlier ledger; failures preserved) | overnight | frontend production build | npm run build | fd87ca2746 | local | 2026-10-21 16:05 | 2026-10-21 16:05 | 0 | passed | — | — | — | — |  | reports/verification/overnight/baseline/logs/build.log
History (earlier ledger; failures preserved) | overnight | e2e normal + recorded judge journeys | npm run e2e -- --reporter=list,junit | fd87ca2746 | recorded data | 2026-10-21 16:05 | 2026-10-21 16:07 | 1 | failed | 27 | 26 | 1 | 0 |  | reports/verification/overnight/baseline/logs/e2e.log | reports/verification/overnight/baseline/junit/e2e.xml
History (earlier ledger; failures preserved) | overnight | e2e public mode (local, synthetic) | npx playwright test --config playwright.public.config.ts --reporter=list,junit | fd87ca2746 | synthetic | 2026-10-21 16:07 | 2026-10-21 16:08 | 0 | passed | 4 | 4 | 0 | 0 |  | reports/verification/overnight/baseline/logs/e2e-public.log | reports/verification/overnight/baseline/junit/e2e-public.xml
History (earlier ledger; failures preserved) | overnight | e2e summary/export | npx playwright test --config playwright.summary.config.ts --reporter=list,junit | fd87ca2746 | local | 2026-10-21 16:08 | 2026-10-21 16:08 | 0 | passed | 3 | 3 | 0 | 0 |  | reports/verification/overnight/baseline/logs/e2e-summary.log | reports/verification/overnight/baseline/junit/e2e-summary.xml
History (earlier ledger; failures preserved) | overnight | README/submission consistency | uv run python scripts/check_submission.py | fd87ca2746 | local | 2026-10-21 16:08 | 2026-10-21 16:08 | 0 | passed | — | — | — | — |  | reports/verification/overnight/baseline/logs/check_submission.log
History (earlier ledger; failures preserved) | overnight | selected-claim consistency (evidence manifest) | uv run python scripts/evidence_manifest.py | fd87ca2746 | local | 2026-10-21 16:08 | 2026-10-21 16:08 | 0 | passed | — | — | — | — |  | reports/verification/overnight/baseline/logs/evidence_manifest.log
History (earlier ledger; failures preserved) | overnight | e2e confirmation: journey.spec.ts review summary (single bounded rerun) | npx playwright test --config playwright.config.ts e2e/journey.spec.ts -g frozen review summary with its cutoff --reporter=list,junit --output=<scratch>/pw-confirm | fd87ca2746 | local | 2026-10-21 16:09 | 2026-10-21 16:09 | 0 | passed | 1 | 1 | 0 | 0 |  | reports/verification/overnight/baseline/logs/e2e-confirm-review-summary.log | reports/verification/overnight/baseline/junit/e2e-confirm.xml
History (earlier ledger; failures preserved) | overnight | parity: registered serving model, default model path, TEMP output | uv run --frozen python scripts/serving_parity.py --output <scratch>/parity_default/serving_parity.json | fd87ca2746 | recorded data | 2026-10-21 16:10 | 2026-10-21 16:10 | 0 | passed | — | — | — | — |  | reports/verification/overnight/logs/parity-default-temp.log
History (earlier ledger; failures preserved) | overnight | parity: --model-dir option on the incumbent model dir (candidate-path smoke) | uv run --frozen python scripts/serving_parity.py --model-dir artifacts/models/lightgbm.dyn-prof-v2.gt-primary-2026-10-13-fusion --output .claude/worktrees/qa/reports/verification/overnight/parity/parity_model_dir_smoke_incumbent.json | fd87ca2746 | recorded data | 2026-10-21 16:10 | 2026-10-21 16:11 | 0 | passed | — | — | — | — |  | reports/verification/overnight/logs/parity-model-dir-incumbent.log
History (earlier ledger; failures preserved) | overnight | python lint (tooling commit) | uv run ruff check . | b391260376 | local | 2026-10-21 16:11 | 2026-10-21 16:11 | 0 | passed | — | — | — | — |  | reports/verification/overnight/tooling/logs/ruff.log
History (earlier ledger; failures preserved) | overnight | backend pytest complete (tooling commit) | uv run pytest -q --junitxml=.claude/worktrees/qa/reports/verification/overnight/tooling/junit/pytest.xml | b391260376 | local | 2026-10-21 16:11 | 2026-10-21 16:13 | 0 | passed | 173 | 172 | 0 | 1 | compared once, from the fusion side (1) | reports/verification/overnight/tooling/logs/pytest.log | reports/verification/overnight/tooling/junit/pytest.xml
History (earlier ledger; failures preserved) | overnight | parity: --model-dir candidate path on the incumbent model (tooling commit) | uv run python scripts/serving_parity.py --model-dir artifacts/models/lightgbm.dyn-prof-v2.gt-primary-2026-10-13-fusion --output .claude/worktrees/qa/reports/verification/overnight/parity/parity_model_dir_smoke_incumbent.json | b391260376 | recorded data | 2026-10-21 16:13 | 2026-10-21 16:14 | 0 | passed | — | — | — | — |  | reports/verification/overnight/tooling/logs/parity-model-dir.log
History (earlier ledger; failures preserved) | overnight | parity: registered serving model, default model path, TEMP output (tooling commit) | uv run python scripts/serving_parity.py --output <scratch>/parity_default/serving_parity.json | b391260376 | recorded data | 2026-10-21 16:14 | 2026-10-21 16:15 | 0 | passed | — | — | — | — |  | reports/verification/overnight/tooling/logs/parity-default-temp.log
History (earlier ledger; failures preserved) | overnight | environment: test_patient_level_data_and_private_notes_are_git_ignored | uv run --no-sync pytest -q -p no:cacheprovider tests/test_replay_api.py::test_patient_level_data_and_private_notes_are_git_ignored | b391260376 | local | 2026-10-21 16:15 |  | 0 | passed | 1 | 1 | 0 | 0
History (earlier ledger; failures preserved) | overnight | tracked-file side effects of the baseline matrix |  | b391260376 | local | 2026-10-21 16:15 |  |  | passed | — | — | — | —
History (earlier ledger; failures preserved) | overnight | python lint (ledger fixes commit) | uv run ruff check . | 622ff0e42b | local | 2026-10-21 17:05 | 2026-10-21 17:05 | 0 | passed | — | — | — | — |  | reports/verification/overnight/tooling/logs/ruff-622ff0e.log
History (earlier ledger; failures preserved) | overnight | backend pytest complete (ledger fixes commit) | uv run pytest -q --junitxml=.claude/worktrees/qa/reports/verification/overnight/tooling/junit/pytest-622ff0e.xml | 622ff0e42b | local | 2026-10-21 17:05 | 2026-10-21 17:07 | 0 | passed | 177 | 176 | 0 | 1 | compared once, from the fusion side (1) | reports/verification/overnight/tooling/logs/pytest-622ff0e.log | reports/verification/overnight/tooling/junit/pytest-622ff0e.xml
History (earlier ledger; failures preserved) | overnight | I python lint | uv run ruff check . | c9b610085a | local | 2026-10-21 22:02 | 2026-10-21 22:02 | 0 | passed | — | — | — | — |  | reports/verification/overnight/integrated/logs/ruff.log
History (earlier ledger; failures preserved) | overnight | I backend pytest (complete) | uv run pytest -q --junitxml=.claude/worktrees/qa/reports/verification/overnight/integrated/junit/pytest.xml | c9b610085a | local | 2026-10-21 22:02 | 2026-10-21 22:04 | 1 | failed | 223 | 219 | 2 | 2 | could not import 'torch': No module named 'torch' (1); compared once, from the fusion side (1) | reports/verification/overnight/integrated/logs/pytest.log | reports/verification/overnight/integrated/junit/pytest.xml
History (earlier ledger; failures preserved) | overnight | I frontend typecheck | npm run typecheck | c9b610085a | local | 2026-10-21 22:04 | 2026-10-21 22:04 | 0 | passed | — | — | — | — |  | reports/verification/overnight/integrated/logs/typecheck.log
History (earlier ledger; failures preserved) | overnight | I frontend unit (vitest) | npm test -- --reporter=default --reporter=junit --outputFile.junit=.claude/worktrees/qa/reports/verification/overnight/integrated/junit/vitest.xml | c9b610085a | local | 2026-10-21 22:04 | 2026-10-21 22:04 | 0 | passed | 133 | 133 | 0 | 0 |  | reports/verification/overnight/integrated/logs/vitest.log | reports/verification/overnight/integrated/junit/vitest.xml
History (earlier ledger; failures preserved) | overnight | I frontend production build | npm run build | c9b610085a | local | 2026-10-21 22:04 | 2026-10-21 22:04 | 0 | passed | — | — | — | — |  | reports/verification/overnight/integrated/logs/build.log
History (earlier ledger; failures preserved) | overnight | I e2e normal + recorded judge journeys | npm run e2e -- --reporter=list,junit --output=<scratch>/pw-out/i-e2e | c9b610085a | recorded data | 2026-10-21 22:04 | 2026-10-21 22:06 | 0 | passed | 29 | 29 | 0 | 0 |  | reports/verification/overnight/integrated/logs/e2e.log | reports/verification/overnight/integrated/junit/e2e.xml
History (earlier ledger; failures preserved) | overnight | I e2e public mode (local, synthetic) | npx playwright test --config playwright.public.config.ts --reporter=list,junit --output=<scratch>/pw-out/i-public | c9b610085a | synthetic | 2026-10-21 22:06 | 2026-10-21 22:06 | 0 | passed | 4 | 4 | 0 | 0 |  | reports/verification/overnight/integrated/logs/e2e-public.log | reports/verification/overnight/integrated/junit/e2e-public.xml
History (earlier ledger; failures preserved) | overnight | I e2e summary/export | npx playwright test --config playwright.summary.config.ts --reporter=list,junit --output=<scratch>/pw-out/i-summary | c9b610085a | local | 2026-10-21 22:06 | 2026-10-21 22:06 | 0 | passed | 3 | 3 | 0 | 0 |  | reports/verification/overnight/integrated/logs/e2e-summary.log | reports/verification/overnight/integrated/junit/e2e-summary.xml
History (earlier ledger; failures preserved) | overnight | I README/submission consistency | uv run python scripts/check_submission.py | c9b610085a | local | 2026-10-21 22:06 | 2026-10-21 22:06 | 0 | passed | — | — | — | — |  | reports/verification/overnight/integrated/logs/check_submission.log
History (earlier ledger; failures preserved) | overnight | I selected-claim consistency (evidence manifest) | uv run python scripts/evidence_manifest.py | c9b610085a | local | 2026-10-21 22:06 | 2026-10-21 22:06 | 0 | passed | — | — | — | — |  | reports/verification/overnight/integrated/logs/evidence_manifest.log
History (earlier ledger; failures preserved) | overnight | I D-053 runner validate (read-only) | uv run python scripts/constrained_study.py validate | c9b610085a | local | 2026-10-21 22:06 | 2026-10-21 22:06 | 0 | passed | — | — | — | — |  | reports/verification/overnight/integrated/logs/constrained_validate.log
History (earlier ledger; failures preserved) | overnight | I parity: incumbent default path, TEMP output | uv run python scripts/serving_parity.py --output <scratch>/i-parity/serving_parity.json | c9b610085a | recorded data | 2026-10-21 22:06 | 2026-10-21 22:07 | 0 | passed | — | — | — | — |  | reports/verification/overnight/integrated/logs/parity-default-temp.log
History (earlier ledger; failures preserved) | overnight | I candidate serving parity |  | b2d76b3b8c | recorded data | 2026-10-21 22:08 |  |  | not run | — | — | — | —
History (earlier ledger; failures preserved) | overnight | I diagnosis: pytest evidence_manifest failures |  | b2d76b3b8c | local | 2026-10-21 22:08 |  |  | failed | — | — | — | —
History (earlier ledger; failures preserved) | overnight | I note: tracked side effects |  | b2d76b3b8c | local | 2026-10-21 22:08 |  |  | passed | — | — | — | —
Author-reported (logs not saved) | fresh clone, 23 Oct | pytest / vitest / e2e on a fresh clone | make test; make test-web; make e2e | unknown | synthetic |  |  | — | author-reported | — | vitest 141; e2e 16 | — | e2e 13 | — | not saved | not saved
Author-reported (no saved CI log) | GitHub Actions CI | ruff, pytest, web jobs | .github/workflows/ci.yml | — | synthetic |  |  | — | author-reported | — | — | — | — | — | not saved | —

## Guardrails
Guardrails — non-numerical behaviour mapped to code and saved tests |  |  |  |  | ← About / navigation
'tested' requires a named test that passed in the saved JUnit of the f073a10 ledger; code alone gives 'implemented'.
Guardrail | Claim | Code / config | Test(s) | Saved result (JUnit, f073a10) | Status | Note
Availability cutoff | Only data available at the issue time is used (labs, meals, wearable, late readings) | ml/glucotwin/features.py; ml/glucotwin/data/canonical.py | tests.test_leakage::test_future_values_cannot_change_earlier_features tests.test_leakage::test_late_logged_meal_is_hidden_until_available tests.test_data_expansion::test_a_reading_released_late_is_invisible_before_its_availability | passed | tested | Saved JUnit of the f073a10 ledger
Native readings only | Interpolated 1-minute rows never count as observations | ml/glucotwin/data/cgmacros.py | tests.test_leakage::test_one_minute_interpolated_rows_do_not_inflate_observations | passed | tested | Saved JUnit of the f073a10 ledger
Matured outcomes for correction | The correction learns only from forecasts whose outcome is observed; never stacks | ml/glucotwin/models/adaptation.py | tests.test_leakage::test_sixty_minute_residual_matures_only_when_target_is_available tests.test_leakage::test_corrections_never_stack_on_corrected_predictions | passed | tested | Saved JUnit of the f073a10 ledger
Bounded correction | ≥ 6 matured outcomes, ±25 mg/dL cap, reset after a 3-h gap | ml/glucotwin/models/adaptation.py; configs/experiment.yaml | tests.test_replay_api::test_personalization_reports_its_bound_and_reset_gap | passed | tested | Saved JUnit of the f073a10 ledger
Abstention | No forecast with insufficient history; nulls, not guesses | ml/glucotwin/features.py | tests.test_replay_api::test_insufficient_history_abstains_with_nulls | passed | tested | Saved JUnit of the f073a10 ledger
Staleness + gap recovery | A real sensor gap marks the forecast stale, the twin abstains, then recovers | ml/glucotwin/features.py; services/api/glucotwin_api/replay.py | e2e: Participant 12: a real sensor gap, abstention without a made-up forecast, then recovery e2e: Participant 12: stale glucose before the gap closes, then abstention | passed | tested | Saved JUnit of the f073a10 ledger
Fusion needs its health record | The profile model abstains until every profile value is available | ml/glucotwin/features.py | tests.test_leakage::test_profile_model_is_gated_until_every_profile_value_is_available | passed | tested | Saved JUnit of the f073a10 ledger
Rewind / revision consistency | Seeking back removes future state; forward replay equals reset + seek | services/api/glucotwin_api/replay.py | tests.test_replay_api::test_seek_backward_removes_future_state_and_invalidates_cursors tests.test_replay_api::test_forward_replay_equals_reset_and_seek | passed | tested | Saved JUnit of the f073a10 ledger
Immutable forecast identity | A served forecast keeps its identity and value (parity with the saved offline forecast) | services/api/glucotwin_api/replay.py; scripts/serving_parity.py | tests.test_v2_preserved::test_model_bundle_reproduces_saved_adaptive_forecasts[gt-primary-2026-10-13-fusion] | passed | tested | Saved JUnit of the f073a10 ledger
Frozen review snapshot | A review summary is one session revision and clock; the copy stays frozen | services/api/glucotwin_api/snapshot.py | tests.test_replay_api::test_review_snapshot_is_one_session_revision_and_clock e2e: Participant 01, Day 9 20:38: a pending warning with a long review note; the copy stays frozen | passed | tested | Saved JUnit of the f073a10 ledger
Warning outcome pending | A warning stays pending until its outcome is observable; gaps censor it | ml/glucotwin/evaluation/alerts.py | tests.test_corrections::test_outcome_stays_pending_until_follow_up_is_observable_and_gaps_censor_it e2e: Participant 01: a warning is pending until its outcome is knowable | passed | tested | Saved JUnit of the f073a10 ledger
Visitor isolation (public mode) | Each browser sees and controls only its own replays | services/api/glucotwin_api/public.py | tests.test_public_demo::test_visitors_cannot_see_or_drive_each_others_sessions e2e: two simultaneous visitors cannot see or control each other's replays | passed | tested | Saved JUnit of the f073a10 ledger
Synthetic labelling | Public mode serves synthetic participants only and labels them | services/api/glucotwin_api/public.py | tests.test_public_demo::test_synthetic_only e2e: a fresh visitor sees the public synthetic demonstration, labelled as such | passed | tested | Saved JUnit of the f073a10 ledger
Test lock / no quiet re-run | Editing a frozen protocol is refused; opened tests need a reason to re-run | ml/glucotwin/research/assess.py | tests.test_research_pipeline::test_editing_a_frozen_protocol_is_refused | passed | tested | Saved JUnit of the f073a10 ledger
Unknown is not zero | Zero onsets/warnings give unavailable metrics; unknown queue counts are never shown as zero | ml/glucotwin/research/metrics.py | tests.test_research_units::test_zero_onsets_and_zero_warnings_give_unavailable_not_perfect_metrics e2e: queue: two replays are counted separately; an unknown count is never shown as zero | passed | tested | Saved JUnit of the f073a10 ledger
No treatment advice / no LLM predictions | No insulin advice, triage or LLM-generated predictions | apps/web; services/api (code inspection) | — | no test cited | implemented | Absence of a feature; code inspection only

## About
GlucoTwin research evidence workbook
Built | 2026-10-24 08:26 UTC from repository HEAD 07a7eac (working tree may differ)
What it is | A judge-readable evidence companion: every public claim → saved source value → visible calculation. Not a new experiment; nothing was fitted or re-opened to build it.
Not | Not clinical validation. A source match shows a number was reported faithfully, not that GlucoTwin helps patients.
Serving test wording | Later recordings of the same 14 people (final test segment) — never 'unseen patients'.
Missing evidence | 'unavailable' / 'not opened' / 'pending' are shown as text, never as 0.
Provenance | Most full artifacts are git-ignored (artifacts/): marked 'local / reproducible-only'. Values are stored here with locators and SHA-256 so they can be checked after re-running the recorded-data path.
Rebuild | uv run --group docs python scripts/build_research_workbook.py ; check: uv run --group docs python scripts/check_research_workbook.py
Method note | docs/submission/GlucoTwin_Research_Workbook_Methodology.md
Sheets
ClaimIndex | Primary view: claims, status, evidence, checks
Sources | Every source value once, with locator and hash
Cohorts | Who each result counts
Forecasts | Errors by evaluation, model and horizon (+ chart)
WarningRanges | Warning denominators, notice shares, range coverage (+ chart)
Studies | D-032, D-043 … D-054 ledger incl. negative and unopened results
Gates | Declared gates on unrounded values
Verification | Test runs, commits, logs; author-reported rows marked
Guardrails | Behaviour → code → saved test
Status meanings
supported | The stated value or fact matches a saved artifact, report or dataset record named in Sources.
supported-arithmetic | Follows by arithmetic from stated definitions (formula shown); no experiment involved.
tested | Behaviour covered by a test that passed in the saved test record (JUnit) named in the row.
implemented | Present in code or configuration only; no saved passing test is cited for the behaviour.
negative-result | A saved study result that did not meet its declared bar; reported as negative, not as a gain.
author-reported | Stated by the team without a saved log, artifact or primary source in the repository.
proposed/illustrative | A proposal, scenario or illustration; not an evidential claim.
pending | Evidence is not on this installation or a primary-source check is still to be done.
unsupported/contradicted | The saved source disagrees with the stated value, or no support exists.